| APPROVALS | DATE | REV | DCN NO. | BY | CHECK | DCC | DATE |
| --- | --- | --- | --- | --- | --- | --- | --- |
| AUTHOR: Mark Lubinski, John Worden | 02/2/09 |  |  |  |  |  |  |
| CHECKED: |  |  |  |  |  |  |  |
| APPROVED: |  |  |  |  |  |  |  |
| DCC RELEASE |  |  |  |  |  |  |  |

Requirements

General

- The oven shall be a front loading machine (not top loading).

- Maximum overall dimensions:  12’x12’ by 10’ in height

- The exterior of the oven shall be insulated to limit thermal hazards and radiation loss.

- The oven will be installed on an existing concrete floor.

- An 8” port shall be provided for possible exhaust provisions

- Unit shall be powered with 480 volt, 3-phase.

- Oven shall be compatible with Class 100 clean room conditions.

- The oven shall meet applicable safety and electrical codes.

- The oven shall fail safe in the event of loss of utilities.

Air bake Chamber

- Minimum chamber load capacity:  8.5’x8.5’ by 36” in height and 3000 lbs in weight (maximum load).

- The bake chamber shall be fabricated of 304SS.

- All materials used in the assembly, such as gaskets, shall be selected to limit particulate.

- No lubrication is permitted within the bake chamber.

- A cart shall be provided to allow loading of 3 plates (1000 lbs each)  at a time.

Oven Controls

- Oven shall have controls for ramp, soak and temp control, and display for internal temperature.

- Soak temp shall be adjustable from 100 deg C to 200 deg C.

- Cycle durations shall be adjustable from a few hours to 7 days.

Heaters and Blower

- The heater shall be sized to allow ramping to 200 deg. C in 8 hours with maximum load.

- There shall be plumbing to allow injection of a Nitrogen purge of 0-200 CFH.

- The recirculation blower shall be of sufficient size to ensure uniform temperature through out bake load.

- The blower lubrication shall be isolated from air flow