## Slide 1: Performance of eLIGO Prototype HAM ISIs and improvements for aLIGO HAM ISIs
Jeffrey Kissel for the SEI team
Graduate, Physics
Louisiana State University
March 19, 2009
LIGO-G0900141

## Slide 2
Advanced LIGO Seismic Isolation
In-vacuum Seismic Isolation
BSC ISI
(LASTI Prototype progress: previous talk)
HAM Chambers
Single Stage ISI
50 at 10 Hz
HAM ISI
(eLIGO Prototype results: this talk)
BSC Chambers
Double Stage ISI
3000 at 1 Hz
Notes: See Fabrice’s talk next!!

## Slide 3
Enhanced LIGO Seismic Isolation
In-vacuum Seismic Isolation
L1 HAM ISI built and installed Feb/Mar ‘08
H1 HAM ISI built and installed Apr/May 2008
L1 HAM ISI commissioning Sep/Oct ’08 and Jan ‘09
H1 HAM ISI commissioning June ‘08 and Feb/Mar ‘09

## Slide 4
X
Y
Z
#1
#2
#3
The HAM ISI
Active Components – Refresher!
Inertial Sensor
Electromagnetic Actuator
Displacement Sensor
Displacement sensors provide alignment and low frequency information
Inertial Sensors (Geophones) provide high frequency information
Electromagnetic actuators provide drive and control of the table at all frequencies
Six sensor clusters mounted symmetrically on the table, one of each of the above in each cluster

## Slide 5
Inertial Sensors
Displacement Sensors
Ground Sensor
Isolation Filters
Hardware
Software
Blend Filters
Damping Filters
HAM ISI “Plant”
Sensor Correction
Super
Sensor
Actuators
The HAM ISI
Active Control Loops – Refresher!
On Board Sensors

## Slide 6
Thanks to R. Mittleman,
C. Gray, and B. Lantz who did all the work!!
Enhanced LIGO Seismic Isolation
RESULTS!!

## Slide 7
Thanks to R. Mittleman,
C. Gray, and B. Lantz who did all the work!!
Enhanced LIGO Seismic Isolation
RESULTS!!

## Slide 8
Thanks to R. Mittleman,
C. Gray, and B. Lantz who did all the work!!
Enhanced LIGO Seismic Isolation
RESULTS!!

## Slide 9
(m/rtHz)
March 10, 2009
H1 at 14 W
33 Hz (in X)
iLIGO Gullwing
n*60 Hz
Power Lines
153 Hz
ISI Blade Spring
110 Hz
The OMC Cage
169 - 200 Hz
The TT
Bounce
362 Hz
ISI Pringle
The eLIGO HAM ISI Noise and Resonance Couplings to DARM
MAYBE
MAYBE

## Slide 10
33 Hz (in X)
iLIGO Gullwing
153 Hz
ISI Blade Spring
110 Hz
The OMC Cage
170 - 200 Hz
The TT
Bounce
362 Hz
ISI Pringle
The eLIGO HAM ISI Noise and Resonance Couplings to DARM
Ideas / Solutions?
Mo-better clamping (Dog Clamps),
Constrained layer damping (Viton)
Constrained layer damping (Viton)
Squeezed film damping (Dash Pot)
Blade springs or thinner suspension wiring
Hrmm…..

## Slide 11
The aLIGO HAM ISI Improving the Performance in Advanced LIGO
The Dreaded Gullwing Resonance:
Build a better support structure, and add HEPI!
Gullwings Crossbeams
Support structure has been redesigned for aLIGO, have already been purchased! (Thanks to Andy Stein @ MIT!)
HEPI will be installed under HAMs, redesigned to use the new crossbeams
Increase f by ~2
Reduce Q by ~10

## Slide 12
The aLIGO HAM ISI Improving the Performance in Advanced LIGO
Found that noise in vertical GS-13s is misconstrued as differential motion >> TILT
TILT couples into horizontal motion degrades performance at low frequencies
Tilt Sensor anyone??
Optimize DISP to GEO (Isolation),
and STS to DISP (Sens. Corr.) blend filters!

## Slide 13
Inertial Sensor
The aLIGO HAM ISI Improving the Performance in Advanced LIGO
The tricked-out SUPER GS-13s
105.2 Hz
Damping ring on vertical configuration’s restoring springs
Stripped down horizontal configuration
NO MORE LOCKING MOTORS!
See Dan Clark’s Poster (G0900128)
Example Features

## Slide 14
The aLIGO HAM ISI Improving the Performance in Advanced LIGO
ISC Group says they need performance better than the requirements between 5 and 20 Hz
We’ll try 6 additional seismometers (L4-Cs) that feed-forward from support stage to suspended stage (Three Horizontal, Three Vertical)
Prototyping fall/winter of 2009 at LASTI
More Feed-forward!

## Slide 15
The HAM ISI So what’s next?
Finish up commissioning at L1 during site shutdown
Find solutions for DARM couplings
Detect gravitational waves
Preliminary and (first half of) Final Reviews finished
Procurement of most parts already purchased or in final stages of approval
Build and commission LASTI prototype in Fall of 2009
Advanced LIGO
Enhanced LIGO

## Slide 17
We have not yet redone the
displacement sensor alignment
since last June
Model (Shown in Purple) shows Y
Spectrum dominated by tilt
Enhanced LIGO Seismic Isolation
Why No Isolation Microseism for X and Y??
AAAHHHH!!!
TILT HORIZONTAL COUPLING!!

## Slide 18
Direction | Gullwing FEA (Hz) | X-Beam FEA (Hz) | Gullwing Meas. From eLIGO (Hz) | X-Beams Meas. @ LASTI (Hz)
X | 6.9 | 11.8 | 7.5 | 9.5
Y | 11.4 | 24.0 | 9.5 | 19
Z | 11.4 | 16.5 | 12 | 16
The HAM ISI Improving the Performance in Advanced LIGO
Installed at LASTI on Dec 18th 2008
Factor of 20 reduction in Q