LIGO Laboratory / LIGO Scientific Collaboration

LIGO-T0900236-v3	    Advanced LIGO UK          4 September 2009

UIM Drive Unit Test Plan

R. M. Cutler, University of Birmingham

Distribution of this document:

Inform aligo_sus

This is an internal working note

of the Advanced LIGO Project, prepared by members of the UK team.

| Institute for Gravitational Research University of Glasgow Phone +44 (0) 141 330 5884 Fax +44 (0) 141 330 6833 E-mail k.strain@physics.gla.ac.uk | School of Physics and Astronomy University of Birmingham Phone +44 (0) 121 414 6447 Fax +44 (0) 121 414 3722 E-mail av@star.sr.bham.ac.uk |
| --- | --- |
| Engineering Department CCLRC Rutherford Appleton Laboratory Phone +44 (0) 1235 445 297 Fax +44 (0) 1235 445 843 E-mail J.Greenhalgh@rl.ac.uk | Department of Physics University of Strathclyde Phone +44 (0) 1411 548 3360 Fax +44 (0) 141 552 2891 E-mail N.Lockerbie@phys.strath.ac.uk |

http://www.ligo.caltech.edu/

http://www.physics.gla.ac.uk/igr/sus/

http://www.sr.bham.ac.uk/research/gravity/rh,d,2.html

http://www.eng-external.rl.ac.uk/advligo/papers_public/ALUK_Homepage.htm

UIM DRIVE

COMPLETED UNIT TEST PLAN

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

Drive Card ID……………

Monitor Card ID ………..

Contents

1. Description

2. Test Equipment

3. Inspection

4. Continuity Checks

5. Test Set Up

6. Power

7. Relay operation

9. Current Monitor Tests

9. Voltage Monitor Tests

10. Corner Frequency Tests

11. Noise Monitor Tests

12. Distortion

13. Final Assembly Tests

1. Description

The UIM Driver Unit will be used to control the mirror position in the Advanced LIGO Gravity wave experiment.

It controls the current in the coil which provides the magnetic force which controls the position of the UIM mirror in a Quad assembly. It works in conjunction with the OSEM coil and position sensor units. One UIM unit controls four OSEMs.

The UIM Coil Drive Unit contains a Coil Drive board and a Monitor board. The Monitor Board monitors the Output voltage, Output Current, RMS Current and Output Noise from the unit.

The UIM Driver Unit also passes the amplified signals from the Photodiodes, which detect the position of the UIM mirror, back to the control electronics without processing them in any way.

Coil Drive Electronics

Monitor

Board

O/Ps to Coils

Coil Demand Signals from AI

Vmon

Imon

I r.m.s

Noise

FIG. 1 UIM Driver Unit Block Diagram

Each UIM Driver Unit consists of four identical differential coil drive channels.

It also contains the monitor board which monitors the output voltage, current, r.m.s current and noise from each channel

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

2. Test Equipment

Power supplies (At least +/- 20v variable, 1A)

Signal generator (capable of delivering 10v peak, 0.1Hz to 10 KHz))

Digital oscilloscope

Analogue oscilloscope

Agilent Dynamic Signal Analyser (or similar)

Low noise Balanced Driver circuit

Relay test box

Record the Models and serial numbers of the test equipment used below.

| Unit (e.g. DVM) | Manufacturer | Model | Serial Number |
| --- | --- | --- | --- |

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

3. Inspection

Remove the lid of the case.

Workmanship

Inspect the general workmanship standard and comment:

Links:

Check that links W3, W4 and W5 are present on each channel of the Drive board. If not, connect them.

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

4. Continuity Checks

Use a multi-meter to check the connections below.

Photodiode outputs

| Pd Out to AA | SIGNAL | DESCRIPTION | Pd in from Sat | OK? |
| --- | --- | --- | --- | --- |
| 1 | PD1P | Photodiode A+ | 1 |  |
| 2 | PD2P | Photodiode B+ | 2 |  |
| 3 | PD3P | Photodiode C+ | 3 |  |
| 4 | PD4P | Photodiode D+ | 4 |  |
| 5 | 0V |  |  |  |
| 6 | PD1N | Photodiode A- | 14 |  |
| 7 | PD2N | Photodiode B- | 15 |  |
| 8 | PD3N | Photodiode C- | 16 |  |
| 9 | PD4N | Photodiode D- | 17 |  |

- LED Monitors

| LED Mon | SIGNAL | Monitors: | In from Sat | OK? |
| --- | --- | --- | --- | --- |
| 1 | Imon1P | Current Source 1+ | 5 |  |
| 2 | Imon2P | Current Source 2+ | 6 |  |
| 3 | Imon3P | Current Source 3+ | 7 |  |
| 4 | Imon4P | Current Source 4+ | 8 |  |
| 5 | 0V |  |  |  |
| 6 | Imon1N | Current Source 1- | 18 |  |
| 7 | Imon2N | Current Source 2- | 19 |  |
| 8 | Imon3N | Current Source 3- | 20 |  |
| 9 | Imon4N | Current Source 4- | 21 |  |

Power Supply to Satellite box

| In from Sat | SIGNAL | DESCRIPTION | DC in Connector | OK? |
| --- | --- | --- | --- | --- |
| 9 | V+ | +17v Supply | A1 |  |
| 10 | V+ | +17v Supply | A1 |  |
| 11 | V- | -17v Supply | A3 |  |
| 12 | V- | -17v Supply | A3 |  |
| 13 | 0V | Return | A2 |  |
| 22 | 0V | Return | A2 |  |
| 23 | 0V | Return | A2 |  |
| 24 | 0V | Return | A2 |  |
| 25 | 0V | Return | A2 |  |

5. TEST SET UP

Signal

Generator

Differential

Driver

Unit under Test

Oscilloscope

Power Supply

Note:

(1)  Input signal to differential amplifier is generally stated in the tests below.

There is therefore an inherent gain of 2 in the system.

(2) Some signal generators will indicate 1vpk/pk when the output is in fact 1v Peak into the high impedance Differential driver used. The test procedure refers to the actual voltage out of the signal generator.

Connections:

Differential signal inputs to the Drive Input of the unit under test:

Drive Input pins 1, 2, 3, 4 = positive input

Drive Input pins 6, 7, 8, 9 = negative input

Drive Input pin 5 = ground

Power (depending on connector availability)

Pd In from Sat pin 9, 10 = +16.5v			or DC in A1

Pd In from Sat pin 11,12 = -16.5 			or DC in A3

Pd In from Sat pins 22, 23, 24, 25 = 0v		or DC in A2

Coil Drive Outputs

Ch1+ = Coil out to Sat pin 1		Ch1- = Coil out to Sat pin 9

Ch2+ = Coil out to Sat pin 3		Ch2- = Coil out to Sat pin 11

Ch3+ = Coil out to Sat pin 5		Ch3- = Coil out to Sat pin 13

Ch4+ = Coil out to Sat pin 7 		Ch4- = Coil out to Sat pin 15

Voltage, Current and R.M.S monitors

1		Voltage Monitor 4

2		Current Monitor 4

3		R.M.S Current   4

4 		Voltage Monitor 3

5 		Current Monitor 3

6 		R.M.S Current   3

7 		Voltage Monitor 2

8 		Current Monitor 2

9 		R.M.S Current   2

10 		Voltage Monitor 1

11 		Current Monitor 1

12		R.M.S Current   1

13 to 25	0v

Noise Monitor

1	Channel 1 Noise Monitor

2 	Channel 2 Noise Monitor

3 	Channel 3 Noise Monitor

4 	Channel 4 Noise Monitor

5 to 9		0v

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

6. Power

Check the polarity of the wiring from the 3 Pin Power Connector, to each of the boards. Viewed from the back of the unit:

A1 	Left pin	Positive	White wire

A2	Middle pin	RTN		Black wire

A3	Right pin	Negative	Green wire

If this is correct,

Connect power to the unit

Set the supplies to 16.5v

Turn on

Record supply currents:

| + 16.5 supply current (mA) | - 16.5 supply current (mA) |
| --- | --- |

Check that all power LEDs are illuminated.

| LEDs | Plus | Minus |
| --- | --- | --- |
| Front Panel |  |  |
| Rear Panel |  |  |

If the power supplies are correct, proceed to the next section.

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

7. Relay Operation

Connect the test unit to the Binary IO inputs.

Operate switches and check that LEDs are on when the relays are switched on, and off when they are switched off:

| Relay | LED Operation |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Filter 1 |  | Filter 2 |  | Filter 2 |  |
|  | ON | OFF | ON | OFF | ON | OFF |
| Ch1 |  |  |  |  |  |  |
| Ch2 |  |  |  |  |  |  |
| Ch3 |  |  |  |  |  |  |
| Ch4 |  |  |  |  |  |  |

This indicates that the relays are operating, and that the talk back contacts work OK.

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

8. Current Monitor tests

The purpose of this test is to perform a functionality test on the current monitor and RMS circuits.

Connect the 39 ohm dummy loads to the coil drive output. Connect the differential driver to the input, fed by the signal generator.  Set the signal generator to 4v at 100 Hz.

Switch out the filters.

Use the true r.m.s meter to measure the r.m.s voltage between test points TP7 and TP11 on the drive board. Adjust the signal generator unit the until this voltage is 3v r.m.s.

Connect a scope probe to each end of one of the load resistors. Set up the scope to differentially observe the voltage across the load resistor.  Check that it is a sine wave.

Measure and record the output from each current monitors, and each of the RMS circuits. Set the meter to r.m.s for the “current monitor” readings (ac signal), and to dc when measuring the r.m.s outputs (dc signal).

The theoretical output is 0.995v.

| Ch. | Monitor Connector | Parameter | Theoretical Value | Measured Value | Pass/Fail |
| --- | --- | --- | --- | --- | --- |
| 1 | Pin 2 | Current Monitor | 0.95v to 1.05 v r.m.s |  |  |
|  | Pin 1 | RMS Current | 0.95 v to 1.05 v dc |  |  |
| 2 | Pin 5 | Current Monitor | 0.95 v to 1.05 v r.m.s |  |  |
|  | Pin 4 | RMS Current | 0.95 v to 1.05 v dc |  |  |
| 3 | Pin 8 | Current Monitor | 0.95 v to 1.05 v r.m.s |  |  |
|  | Pin 7 | RMS Current | 0.95 v to 1.05 v dc |  |  |
| 4 | Pin 11 | Current Monitor | 0.95 v to 1.05 v r.m.s |  |  |
|  | Pin 10 | RMS Current | 0.95 v to 1.05 v dc |  |  |

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

9. Voltage Monitor tests

The purpose of this test is to verify and calibrate the Voltage Monitor circuit on each channel.

Switch all filters out. Remove the dummy loads and make differential voltage output measurements on the coil drive outputs TP7 and TP11 at 100 Hz. Adjust the signal generator to give a voltage to 3v r.m.s on the coil drive outputs.

Record the peak voltage on each Voltage Monitor pin, and check against the theoretical figure.

| Channel | Coil Drive Output pins | Voltage Monitor socket Pin | Monitor output? | Expected value | OK? |
| --- | --- | --- | --- | --- | --- |
| 1 | Pins 1,9 | Pin 3 |  | 0.95v to 1.05v |  |
| 2 | Pins 3,11 | Pin 6 |  | 0.95v to 1.05v |  |
| 3 | Pins 5,13 | Pin 9 |  | 0.95v to 1.05v |  |
| 4 | Pins 7, 15 | Pin 12 |  | 0.95v to 1.05v |  |

Measurements may be used using a 25 way break out box, or a 25 way connector on the V, I and r.m.s connector.

10. Corner frequency tests

The purpose of this test is to verify that the frequency response of each filter stage of each channel is correct. There are two possible ways of performing this test depending on available equipment. Either 8.1 or 8.2 should be conducted (it is not necessary to do both).

10.1 Using a Signal Generator

Apply a signal to the input, amplitude 1v peak, Frequency 100Hz via the differential driver, to the Drive Input.

.

Switch out all the filter stages. Measure and record the peak differential output between the outputs of each channel, using two oscilloscope channels in differential mode.

|  | Coil Drive output pins | Pk/pk Output | Specification | Pass/Fail |
| --- | --- | --- | --- | --- |
| Ch1 | 1 and 9 |  | 4.8v to 5v |  |
| Ch2 | 3 and 11 |  | 4.8v to 5v |  |
| Ch3 | 5 and 13 |  | 4.8v to 5v |  |
| Ch4 | 7 and 15 |  | 4.8v to 5v |  |

Switch in each filter in turn and measure and record the peak output at 1Hz, 10 Hz, and 1 KHz. 			     1Hz

|  | Coil Drive output pins | F1 | F2 | F3 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- | --- | --- |
| Ch1 | 1 and 9 |  |  |  | 3.3v to 3.7v |  |
| Ch2 | 3 and 11 |  |  |  | 3.3v to 3.7v |  |
| Ch3 | 5 and 13 |  |  |  | 3.3v to 3.7v |  |
| Ch4 | 7 and 15 |  |  |  | 3.3v to 3.7v |  |

10Hz

|  | Coil Drive output pins | F1 | F2 | F3 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- | --- | --- |
| Ch1 | 1 and 9 |  |  |  | 0.48 to 0.75v |  |
| Ch2 | 3 and 11 |  |  |  | 0.48 to 0.75v |  |
| Ch3 | 5 and 13 |  |  |  | 0.48 to 0.75v |  |
| Ch4 | 7 and 15 |  |  |  | 0.48 to 0.75v |  |

1 KHz

|  | Coil Drive output pins | F1 | F2 | F3 | Specification | Pass/Fail |
| --- | --- | --- | --- | --- | --- | --- |
| Ch1 | 1 and 9 |  |  |  | 0.4v to 0.5v |  |
| Ch2 | 3 and 11 |  |  |  | 0.4v to 0.5v |  |
| Ch3 | 5 and 13 |  |  |  | 0.4v to 0.5v |  |
| Ch4 | 7 and 15 |  |  |  | 0.4v to 0.5v |  |

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

10.2. Using the Dynamic Signal Analyser

With all filters switched in, measure the frequency response of each channel in turn between 0.1 Hz and 1KHz. If a fast turn around is required, limit the measurement to the frequency range to between 1 Hz and 1KHz. Measure the gain at the spot frequencies below and record them.

Connect the 39 ohm loads across each coil output to simulate the coils.

Channel 1

| Frequency | Gain (dB) | Expected Gain (dB) | Pass/Fail |
| --- | --- | --- | --- |
| 0.1 Hz |  | 6.6 to 8.7 |  |
| 1Hz |  | -0.2 to -2 |  |
| 10Hz |  | -41 to -52 |  |
| 100Hz |  | -51 to -58 |  |
| 1KHz |  | -51 to -58 |  |

Channel 2

| Frequency | Gain (dB) | Expected Gain (dB) | Pass/Fail |
| --- | --- | --- | --- |
| 0.1 Hz |  | 6.6 to 8.7 |  |
| 1Hz |  | -0.2 to -2 |  |
| 10Hz |  | -41 to -52 |  |
| 100Hz |  | -51 to -58 |  |
| 1KHz |  | -51 to -58 |  |

Channel 3

| Frequency | Gain (dB) | Expected Gain (dB) | Pass/Fail |
| --- | --- | --- | --- |
| 0.1 Hz |  | 6.6 to 8.7 |  |
| 1Hz |  | -0.2 to -2 |  |
| 10Hz |  | -41 to -52 |  |
| 100Hz |  | -51 to -58 |  |
| 1KHz |  | -51 to -58 |  |

Channel 4

| Frequency | Gain (dB) | Expected Gain | Pass/Fail |
| --- | --- | --- | --- |
| 0.1 Hz |  | 6.6 to 8.7 |  |
| 1Hz |  | -0.2 to -2 |  |
| 10Hz |  | -41 to -52 |  |
| 100Hz |  | -51 to -58 |  |
| 1KHz |  | -51 to -58 |  |

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

11. Noise Tests

Ensure the filter links W2, W3, W4 and W5 are present on each channel.

Connect the filter test box, and switch in all filters.

Switch it out of Test Mode

Use the HP 35670A Dynamic Signal Analyser.

Connect a shorting plug to the demand input to short all positive and negative demands together and to 0v. Connect 39 Ohm loads to the outputs.

Switch the filters in.

Use Stuart Aston’s noise measurement set up, loaded from disc.

Measure the noise output from each channel in turn at the amplifier outputs. The Low Pass filter on the SR650 may be used to reduce mains interference, to prevent the Signal Analyser from overloading. Ideally the filter corner frequency should be set to 3 KHz, and readings should be made when the ambient noise is low. Set the amplifier gain to 1000, and check that the overload light is not on before each measurement.

Unit under Test

Power Supply

HP 35670A Signal Analyser

Demand

Inputs

0V

SR560

|  | Spec in dB V/√Hz | Measured @ 10Hz | -60dB = |
| --- | --- | --- | --- |
| Ch1 | -152.6 |  |  |
| Ch2 | -152.6 |  |  |
| Ch3 | -152.6 |  |  |
| Ch4 | -152.6 |  |  |

The specification at 10 Hz is a noise current of 3 pA per root Hz. The total load resistance is 7.8k. The permitted noise voltage from the amplifiers is therefore 23.4 nA/ /√Hz or -152.6 dB. The SR560 is set to a gain of 1000 (60dB) so 60dB needs to be subtracted from the reading.

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

12. Distortion

No filters. Increase input voltage to 10v peak, f = 1KHz. Dummy 39 Ohm loads. Check the differential voltage across each load for distortion with an analogue oscilloscope.

|  | Distortion Free? |
| --- | --- |
| Ch1 |  |
| Ch2 |  |
| Ch3 |  |
| Ch4 |  |

Unit……………………………………Serial No ………………………………

Test Engineer ………

Date …………………

13. Final Assembly Tests

1. Remove the lid of the box.

2. Unplug all external connections.

3. Check that the 9mm pillars are in place in the corners of the Monitor Board towards the centre of the box.

4. Check that all internal connectors are firmly mated.

5. Tighten the screw-locks holding all the external connectors.

6. Check that all the LEDs are nicely centred.

7. Check that links W3, W4 and W5 are in place.

8. Check that the boards are labelled with their Drawing Number, Issue Number, and serial number. Record below:

| UoB box ID |  |
| --- | --- |
| Driver board ID |  |
| Driver board Drawing No/Issue No |  |
| Driver board Serial Number |  |
| Monitor board ID |  |
| Monitor board Drawing No/Issue No |  |
| Monitor board Serial Number |  |

9. Check the security of any modification wires.

10. Visually inspect.

11. Put the lid on and fasten all screws,

Check all external screws for tightness.