| APPROVALS | DATE | REV | DCN NO. | BY | CHECK | DCC | DATE |
| --- | --- | --- | --- | --- | --- | --- | --- |
| AUTHOR: |  |  |  |  |  |  |  |
| CHECKED: |  |  |  |  |  |  |  |
| APPROVED: |  |  |  |  |  |  |  |
| DCC RELEASE |  |  |  |  |  |  |  |

SPECIFICATION TABLE OF CONTENTS

1.0	Purpose

2.0	General

3.0	Responsibilities

4.0	Procedure

5.0	Required Documentation

# 1.0	PURPOSE

This specification covers the procedure to be used for black light inspections of cleaned LIGO Vacuum Components. The Seller shall incorporate these requirements in their procedures.

# 2.0	GENERAL

2.1	Cleaning shall be performed by an engineer, technician or shop worker familiar with this procedure and the manual for use of the ultraviolet lamp.

2.2	Inspection shall be done in a darkened area.

2.3	A 100 watt ultraviolet lamp with a wavelength of 365 nm shall be used.

# 3.0	RESPONSIBILITIES

3.1	Cleaning shall be performed by an engineer, technician or shop worker familiar with this procedure and the manual for use of the ultraviolet lamp.

# 4.0	PROCEDURE

4.1	Turn on the lamp and allow it to warm up for 3-5 minutes.

4.2	The operator shall enter the darkened area and allow 2-3 minutes for his eyes to adapt to the low light level.

4.3	Inspect the cleaned surfaces, holding the lamp 8-12 inches from the surface.  Be careful to distinguish between ultraviolet fluorescence and reflected purple visible light.

4.4	Record the description and location of any detected contamination and issue a Discrepancy Report.

# 5.0	Required Documentation

Inspection results shall be recorded on the component cleaning data sheet.