## how to use this workbook
How to use this workbook:
1) Get an E-number for this test report from https://dcc.ligo.org
a) Label the document "Optical Contamination Test Results: <name>", where <name> is a descriptive name of the article under test
b) For the E-<number>-X0 version DCC entry, give a COMPLETE description of the article(s) under test in the Abstract (not the "Changes and Notes"), including manufacturer/source, material grade or assembly/device model number, and quantity/amount. For example, it is not sufficient to say "photodiode", but rather state the manufacturer's name and model number, such as "Vishay TSTS7100 IR-LED". Likewise "PEEK" is not a sufficient material descriptor, but "Victrex PEEK grade 450GL" is a complete description.
c) For the E-<number>-X0 version DCC entry, give a URL link to the location where intermediate test results will be posted in the "Notes and Changes" field.
2) Rename this template workbook to "E<number>-v1 <name>.xls"
3) Fill in only the light yellow colored cells in the sheet labelled "scaled Loss rate".
You must decide the appropriate units to characterize the material, or assembly or device under test. For example, for a device or assembly the appropriate units is likely "each", whereas for a bulk material (say PEEK) the appropriate unit is likely surface area (cm^2). For a lubricant, such as Krytox, the appropriate units may be mass, or volume.
4) Fill in only the light yellow colored cells in the sheets labelled "d1", "d2", "d3", etc.
Each sheet is a data set taken at a particular day and time. In the first data set sheet (d1), the start date of the test and the cavity number is entered. Data for the TEM 00-01 beat frequency as a function of power and the ring down data are entered for each data set. The data is plotted and fitted automatically. The sheet labelled "data" summarizes the absorption and total loss measurements as a function of time in a tabular listing. Sheets "absorption" and "total loss" summarize the absortion and total loss in plots and a least squares ft, as the data is entered. The "scaled loss rate" sheet summarizes the results scaled to the LIGO Observatory for the intended use and is updated automatically as data is entered.
5) The test is terminated when a determination is made that the material/assembly/device under test has either passed or failed. At this point, use the "Convert to Adobe PDF" tool to create a file with just the spredsheets labelled "scaled loss rate", "absorption" and "total loss" which is named "E<number>-v1 <name>.pdf".
6) Select "Create E<number>-v1" in the DCC entry for this document and upload the pdf file "E<number>-v1 <name>.pdf" as the "Main file in document" and this Excel file "E<number>-v1 <name>.xls" as an "Other file" (not Main). Then in either the abstract or the "Notes and Changes" field indicate the test results (passed or failed).
7) File any other documents, which help to define the material/assembly/device which was tested, as "Other files" under the DCC entry for E<number>-v1, such as photos or descriptions of the test (if it was a unique test, such as taken at elevated temperatures).
8) File proprietary data (data not owned by LIGO Lab, even if posted to the web) about the material/assembly/device such as MSDS sheets, specifications, etc. This data should be filed under a separate E-number and then linked from the E-number which corresponds to the optical contamination cavity test results. Even if the MSDS, specifications, etc. are already on the web, they should be filed in the DCC, in case the material/assembly/device is discontinued (in which case the manufacturer often removes this data from their web sites.)

## scaled loss rate
LIGO E-Document Number: |  |  | E1000192
Sample Test:
| Material under test: |  | Q3000 Quad IR Photodiodes (1064nm)
| units |  |  | 6 | each
| absorption | 0.033993 | ± | 0.37528 | ppm/yr | 1 sigma
| scatter | 6.9748 | ± | 5.0595 | ppm/yr | 1 sigma
| max. normalized absorption |  |  | 0.13076 | ppm/yr/unit | 2 sigma
| max. normalized scatter |  |  | 2.849 | ppm/yr/unit | 2 sigma
| test turbopump speed (liter/s) |  |  | 24.3916 | torr/liter/sec
Scaled to LIGO:
| LIGO Vacuum Volume |  |  | Vertex | LHO Diagonal | End | Comments
| Quantity (units) |  |  | 8 | 8 | 2 | 2 in HAM1/7 and 6 in HAM6/12. 2 in vertex & diagonal, but might open HAM6/12 to verter/diagonal
| LIGO ion pumping speed (liter/s) |  |  | 6800 | 6800 | 1700 | see E0900398 or PSI V049-1-078 for pump rates
| pumping speed ratio (test/LIGO) |  |  | 0.003587 | 0.003587 | 0.014348 | does not include cryo-pump and effective pumping from the Beam Tube
| max. absorption (ppm/yr) |  |  | 0.0037522 | 0.0037522 | 0.0037522 | * Limit is < 0.02 ppm/yr for a single source
| max. scatter (ppm/yr) |  |  | 0.081754 | 0.081754 | 0.081754 | * Limit is < 0.2 ppm/yr for a single source
* The overall limit on contamination loss on optics for AdL is < 0.5 ppm/yr absorption and < 4 ppm/yr scatter from all sources, per Table 4 of the COC Design Requirements Document (T000127-v1). It is assumed that ~20 sources could contribute.
The printing on the side of the C30845 canswas removed by and electric soft rubber eraser and sonic cleaned for two minutes in isopropyl alcohol then wrensed in
Methanol for 5 minutes and then the tops were cut off before RGA tests were done. Then put into the contamination test cavity for final testing.
|  |  |  | 2.75 | in, dia, flex hose
|  |  |  | 36 | in, length, flex hose
|  |  |  | 6.985 | cm, dia
|  |  |  | 91.44 | cm, length
|  |  |  | 463.669 | C1, aperture, liter/sec
|  |  |  | 45.0971 | C2, flex hose, liter/sec
|  |  |  | 41.0997 | Ct, liter/sec
|  |  |  | 60 | S, liter/sec [Varian Dry TPS-compact with TV81M]
|  |  |  | 24.3916 | Seff, liter/sec

## absorption
LIGO E-Document Number: |  | E1000192
Test Material/Assy./Device: |  | Q3000 Quad IR Photodiodes (1064nm)
Absorption fitting
| Slope | 3.8805e-06 | 1.1068 | Y-intercept
| Standard Error | 4.284e-05 | 0.021509 | Standard error
| r2 | 0.00091083 | 0.02872 | sey
| F | 0.0082049 | 9 | df
| ssreg | 6.7677e-06 | 0.0074235 | ssresid
| Absorption change rate (ppm/yr) |  | ± sigma (ppm/yr)
| 0.033993 |  | 0.37528

## total loss
LIGO E-Document Number: |  | E1000192
Test Material/Assy./Device: |  | Q3000 Quad IR Photodiodes (1064nm)
Total loss fitting
| Slope | 0.00079621 | 146.2935 | Y-intercept
| Standard Error | 0.00057757 | 0.28998 | Standard error
| r2 | 0.17434 | 0.3872 | sey
| F | 1.9004 | 9 | df
| ssreg | 0.28492 | 1.3493 | ssresid
| Total loss change rate (ppm/yr) |  | ± sigma (ppm/yr)
| 6.9748 |  | 5.0595

## data
| Time (hr) | Absorption (ppm) | Total Loss (ppm) | Q3000 Quad IR Photodiodes (1064nm)
1 | 131 | 1.0758 | 146.1988
2 | 182 | 1.1361 | 146.8045
3 | 272.5 | 1.0698 | 146.5944
4 | 346 | 1.1402 | 146.7999
5 | 440 | 1.1142 | 146.0791
6 | 466.5 | 1.1028 | 146.8959
7 | 512.5 | 1.1465 | 146.6794
8 | 586 | 1.1274 | 146.627
9 | 634.5 | 1.0746 | 146.691
10 | 683.5 | 1.1068 | 146.3008
11 | 801 | 1.1 | 147.5832
12 | #N/A | #N/A | #N/A
13 | #N/A | #N/A | #N/A
14 | #N/A | #N/A | #N/A
15 | #N/A | #N/A | #N/A
16 | #N/A | #N/A | #N/A
17 | #N/A | #N/A | #N/A
18 | #N/A | #N/A | #N/A
19 | #N/A | #N/A | #N/A
20 | #N/A | #N/A | #N/A
21 | #N/A | #N/A | #N/A
22 | #N/A | #N/A | #N/A
23 | #N/A | #N/A | #N/A
24 | #N/A | #N/A | #N/A
25 | #N/A | #N/A | #N/A
26 | #N/A | #N/A | #N/A
27 | #N/A | #N/A | #N/A
28 | #N/A | #N/A | #N/A
29 | #N/A | #N/A | #N/A
30 | #N/A | #N/A | #N/A
31 | #N/A | #N/A | #N/A
32 | #N/A | #N/A | #N/A
33 | #N/A | #N/A | #N/A
34 | #N/A | #N/A | #N/A
35 | #N/A | #N/A | #N/A
36 | #N/A | #N/A | #N/A
37 | #N/A | #N/A | #N/A
38 | #N/A | #N/A | #N/A
39 | #N/A | #N/A | #N/A
40 | #N/A | #N/A | #N/A
41 | #N/A | #N/A | #N/A
42 | #N/A | #N/A | #N/A
43 | #N/A | #N/A | #N/A
44 | #N/A | #N/A | #N/A
45 | #N/A | #N/A | #N/A
46 | #N/A | #N/A | #N/A
47 | #N/A | #N/A | #N/A
48 | #N/A | #N/A | #N/A
49 | #N/A | #N/A | #N/A
50 | #N/A | #N/A | #N/A

## d1
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 | 2010-06-02 | 11 | 5 | 131 | 2.43
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
24 | 6.6e-09 | 25 | 40.2746
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
64 | 0.438 | 0.156 | 74.477 | 2.28e-05 | 1.76e-08
70 | 0.434 | 0.147 | 74.479
74 | 0.432 | 0.135 | 74.482
78 | 0.399 | 0.118 | 74.487
82 | 0.373 | 0.098 | 74.493
86 | 0.35 | 0.077 | 74.499
90 | 0.32 | 0.0517 | 74.505
94 | 0.295 | 0.0328 | 74.511
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.0758 |  | 146.1988

## d2
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 | 2010-06-04 | 14 | 7 | 182 | 2.43
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
24.5 | 6.4e-09 | 25 | 40.2631
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
64 | 0.422 | 0.156 | 74.4778 | 2.2706e-05 | 1.65e-08
70 | 0.41 | 0.142 | 74.4808
74 | 0.395 | 0.135 | 74.4845
78 | 0.38 | 0.119 | 74.489
82 | 0.365 | 0.1 | 74.495
86 | 0.345 | 0.078 | 74.501
90 | 0.325 | 0.054 | 74.5078
94 | 0.308 | 0.034 | 74.513
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.1361 |  | 146.8045

## d3
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 | 2010-06-08 | 8.5 | 11 | 272.5 | 2.43
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
23.3 | 6.3e-09 | 25.1 | 40.2572
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
64 | 0.437 | 0.154 | 74.476 | 2.2738e-05 | 9.11e-09
70 | 0.43 | 0.147 | 74.478
74 | 0.415 | 0.134 | 74.482
78 | 0.396 | 0.117 | 74.486
82 | 0.374 | 0.098 | 74.492
86 | 0.348 | 0.0743 | 74.498
90 | 0.321 | 0.0527 | 74.504
94 | 0.301 | 0.035 | 74.509
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.0698 |  | 146.5944

## d4
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 | 2010-06-11 | 10 | 14 | 346 | 2.43
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
23.6 | 6.1e-09 | 25 | 40.2411
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
64 | 0.417 | 0.157 | 74.474 | 2.2707e-05 | 8.2e-09
70 | 0.411 | 0.149 | 74.476
74 | 0.403 | 0.135 | 74.481
78 | 0.384 | 0.119 | 74.486
82 | 0.366 | 0.099 | 74.492
86 | 0.346 | 0.079 | 74.498
90 | 0.323 | 0.0536 | 74.504
94 | 0.306 | 0.0348 | 74.51
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.1402 |  | 146.7999

## d5
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 | 2010-06-15 | 8 | 18 | 440 | 2.43
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
23.7 | 6.1e-09 | 25.1 | 40.237
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
64 | 0.416 | 0.156 | 74.476 | 2.2819e-05 | 7.5e-09
70 | 0.409 | 0.147 | 74.48
74 | 0.398 | 0.136 | 74.483
78 | 0.382 | 0.119 | 74.488
82 | 0.362 | 0.099 | 74.495
86 | 0.339 | 0.077 | 74.501
90 | 0.322 | 0.0538 | 74.506
94 | 0.3 | 0.0355 | 74.511
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.1142 |  | 146.0791

## d6
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 | 2010-06-16 | 10.5 | 19 | 466.5 | 2.43
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
23.4 | 6.1e-09 | 25 | 40.237
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
64 | 0.426 | 0.156 | 74.475 | 2.2692e-05 | 7.96e-09
70 | 0.415 | 0.147 | 74.478
74 | 0.404 | 0.135 | 74.481
78 | 0.386 | 0.118 | 74.486
82 | 0.363 | 0.098 | 74.492
86 | 0.34 | 0.077 | 74.498
90 | 0.315 | 0.0533 | 74.504
94 | 0.297 | 0.0339 | 74.51
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.1028 |  | 146.8959

## d7
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 | 2010-06-18 | 8.5 | 21 | 512.5 | 2.43
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
23.2 | 6.1e-09 | 25 | 40.237
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
64 | 0.436 | 0.153 | 74.474 | 2.2725e-05 | 9.64e-09
70 | 0.424 | 0.146 | 74.477
74 | 0.411 | 0.135 | 74.481
78 | 0.394 | 0.118 | 74.486
82 | 0.372 | 0.099 | 74.492
86 | 0.344 | 0.077 | 74.498
90 | 0.319 | 0.0527 | 74.504
94 | 0.297 | 0.0344 | 74.51
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.1465 |  | 146.6794

## d8
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 | 2010-06-21 | 10 | 24 | 586 | 2.43
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
23.4 | 6e-09 | 25 | 40.237
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
64 | 0.416 | 0.154 | 74.476 | 2.2733e-05 | 9.21e-09
70 | 0.411 | 0.146 | 74.477
74 | 0.4 | 0.134 | 74.482
78 | 0.38 | 0.118 | 74.487
82 | 0.36 | 0.098 | 74.493
86 | 0.337 | 0.077 | 74.5
90 | 0.317 | 0.0526 | 74.505
94 | 0.297 | 0.0339 | 74.51
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.1274 |  | 146.627

## d9
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 | 2010-06-23 | 10.5 | 26 | 634.5 | 2.43
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
24 | 6e-09 | 25 | 40.2247
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
64 | 0.436 | 0.151 | 74.477 | 2.2724e-05 | 1.07e-08
70 | 0.427 | 0.142 | 74.48
74 | 0.414 | 0.13 | 74.483
78 | 0.395 | 0.115 | 74.488
82 | 0.377 | 0.095 | 74.493
86 | 0.351 | 0.074 | 74.499
90 | 0.326 | 0.0513 | 74.505
94 | 0.304 | 0.0324 | 74.51
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.0746 |  | 146.691

## d10
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 | 2010-06-25 | 11.5 | 28 | 683.5 | 2.43
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
24.3 | 6e-09 | 25 | 40.3332
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
66 | 0.472 | 0.163 | 74.477 | 2.2784e-05 | 2.73e-08
70 | 0.461 | 0.153 | 74.48
74 | 0.447 | 0.142 | 74.482
78 | 0.425 | 0.124 | 74.489
82 | 0.404 | 0.103 | 74.494
86 | 0.38 | 0.08 | 74.501
90 | 0.354 | 0.0545 | 74.5078
94 | 0.332 | 0.0312 | 74.5145
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.1068 |  | 146.3008

## d11
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 | 2010-06-30 | 9 | 33 | 801 | 2.43
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
28.1 | 5.7e-09 | 26.3 | 40.2361
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
64 | 0.618 | 0.165 | 74.477 | 2.2586e-05 | 1.26e-08
70 | 0.6 | 0.156 | 74.481
74 | 0.584 | 0.143 | 74.485
78 | 0.561 | 0.125 | 74.49
82 | 0.528 | 0.103 | 74.497
86 | 0.495 | 0.082 | 74.502
90 | 0.464 | 0.0557 | 74.51
94 | 0.429 | 0.0348 | 74.514
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.1 |  | 147.5832

## d12
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 |  |  |  | 0
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | #DIV/0!

## d13
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 |  |  |  | 0
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | #DIV/0!

## d14
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 |  |  | -40326 | -967824
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | #DIV/0!

## d15
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 |  |  | -40326 | -967824
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | #DIV/0!

## d16
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 |  |  | -40326 | -967824
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | #DIV/0!

## d17
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 |  |  | -40326 | -967824
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | #DIV/0!

## d18
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 |  |  | -40326 | -967824
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | #DIV/0!

## d19
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 |  |  | -40326 | -967824
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | #DIV/0!

## d20
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 |  |  | -40326 | -967824
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | #DIV/0!

## d21
Cavity No | Sample | Comment
2 | Q3000 Quad IR Photodiodes (1064nm)
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-05-28 |  |  | -40326 | -967824
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | #VALUE! |  | #DIV/0!