LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1000330 V2	                 LIGO                              Oct 26th, 2014

aLIGO HAM-ISI, Pre-integration Test Report, Phase I, LLO India #1 (post-assembly, before storage)

E1000330 – V2

Adrien Le Roux, Hugo Paris (computation of missing LZMP only)

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

Table of contents:

Introduction	4

I.	Pre-Assembly Testing	5

	Step 1: Position Sensors	5

	Step 2: GS13	7

	Step 2.1 – Horizontal GS-13s	7

Step 2.2 – Vertical GS-13s	9

	Step 3: Actuators	11

II.	Tests to be performed during assembly	12

	Step 1: Inventory (E1000052)	12

	Step 2: Check torques on all bolts	12

	Step 3: Check gaps under Support Posts	12

	Step 4: Pitchfork/Boxwork flatness before Optical Table install	13

	Step 5: Blade spring profile	13

	Step 6: Gap checks on actuators-after installation on Stage 1	14

	Step 7: Check level of Stage 0	14

	Step 8: Check level of Stage 1 Optical Table	15

	Step 9: Mass budget	16

	Step 10: Shim thickness	17

	Step 11: Lockers adjustment	17

III.	Tests to be performed after assembly	18

	Step 1 - Electronics Inventory	18

	Step 2 - Set up sensors gap	18

	Step 3 - Measure the Sensor gap	19

	Step 4 - Check Sensor gaps after the platform release	19

	Step 5 – Performance of the limiter	20

	Step 5.1 - Test Nº1 - Push “in the general coordinates”	20

	Step 5.2 - Test Nº2 – Push “locally”	20

	Step 6 - Position Sensors unlocked/locked Power Spectrum	21

	Step 7 - GS13 power spectrum -tabled tilted	23

	Step 8- GS13 pressure readout	25

	Step 9 - Coil Driver, cabling and resistance check	26

	Step 10 - Actuators Sign and range of motion (Local drive)	27

	Step 11 - Vertical Sensor Calibration	28

	Step 12 - Vertical Spring Constant	29

	Step 13 - Static Testing (Tests in the local basis)	29

	Step 14 - Linearity test	31

	Step 15 - Cartesian Basis Static Testing	33

	Step 16- Frequency response	34

	Step 16.1 - Local to local measurements	34

	Step 16.2 - Cartesian to Cartesian measurements	37

	Step 17 - Transfer function comparison with Reference	38

	Step 17.1 - Local to local - Comparison with Reference	38

	Step 17.2 - Cartesian to Cartesian - Comparison with Reference	42

	Step 18 - Lower Zero Moment Plane	43

	Step 19 - Damping loops	45

	Step 19.1 - Transfer functions - Simulation	45

	Step 19.2 - Powerspectra – Experimental	46

Conclusion	1

Introduction

This first Unit for LIGO India was built in fall 2013, and was tested in October/November 2013, just after we got done building BSC-ISI at LLO.

The procedure document used to perform this test is:

- E1000309–V9 - aLIGO HAM-ISI, Pre-Integration Testing Procedure, Phase I (post assembly, before storage)

Other useful information can be found in:

-  E1000300 - HAM-ISI LLO test stand: software and electronic check

# Pre-Assembly Testing

## Step 1: Position Sensors

| S/N sensor | S/N board | ADE Gap Standoff(mm) | Location on the Jig | Gap Standoff on Jig(mm/in) | Voltage before zeroing | Voltage after zeroing. Prebake | Voltage after zeroing. Post bake |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 11982 | NR | NR | NR | NR | NR | NR | NR |
| 11995 | NR | NR | NR | NR | NR | NR | NR |
| 12010 | NR | NR | NR | NR | NR | NR | NR |
| 12003 | NR | NR | NR | NR | NR | NR | NR |
| 11993 | NR | NR | NR | NR | NR | NR | NR |
| 12037 ( 12058 | NR | NR | NR | NR | NR | NR | NR |

NR: not recorded

Note: CPS V3 12037 was replaced by 12058 because its behavior was different than the other verticals in the TFs.

Sensors noise spectra measured before baking:

| Figure 1: H1 and V1 sensor noise | Figure 2: H2 and V2 sensor noise |
| --- | --- |

Figure 3: H3 (12010) and V3 (12037 & 12058) sensor noise

Acceptance Criteria:

- Power spectrum magnitudes must be lower than:

  - 9.e-10 m/√Hz at 0.1Hz

  - 6.e-10 m/√Hz at 1Hz

Issues/difficulties/comments regarding this test:

Test result:	Passed:   X   	Failed:         .

## Step 2: GS13

All the data related to GS-13 post podding testing can be found in the SVN at : SeismicSVN\seismic\Common\Data\aLIGO_GS13_TestData\PostMod_TestResults_PDFs

E1000058 spreadsheet provides the status of each individual GS-13 at LLO site during aLIGO HAM assembly

Data files in SVN at:

/ligo/svncommon/SeiSVN/seismic/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_RawASCII

Scripts files for processing and plotting in SVN at:

/ligo/svncommon/SeiSVN/seismic/Common/MatlabTools

- gs13qatest.m

Figures in SVN at:

/ligo/svncommon/SeiSVN/seismic/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_PDFs

## Step 2.1 – Horizontal GS-13s

Huddle testing

Figure 4: Huddle testing of Horiz GS-13 796, 775, and 807 after aLIGO modifications

## Step 2.2 – Vertical GS-13s

Huddle testing

Figure 5: Huddle testing of Vert GS-13 769, 702 and 763 after aLIGO modifications

Driven testing

Figure 6: Driven Transfer Function of Vert GS-13 769, 702 and 763 after aLIGO modifications

Acceptance Criteria:

- GS13 should have been already modified and tested. GS-13 Inspection/Pod Assembly (D047810). Checklist is defined in F090070-v6

- Resonant frequency at 1 Hz (huddle testing)

- No spring resonance on vertical driven tests

Test result:	Passed:   X  	Failed:         .

## Step 3: Actuators

Actuator data can be found at: T0900564. Actuator inventory is made at Section II – Step 1.

| Actuator Serial #: L133 Operator Name: Gordon, Matt Date: 4/12/2010 Time: 2:51 PM Actuator Coil Resistance: 6.43 Ohms, PASS Ambient Temperature: 73.3 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.529 Y Travel Limit (inches): 0.205 Z Travel Limit (inches): 0.505 | Actuator Serial #: L069 Operator Name: Gordon, Matt Date: 11/23/2009 Time: 9:41 AM Actuator Coil Resistance: 6.29 Ohms, PASS Ambient Temperature: 69.8 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.537 Y Travel Limit (inches): 0.205 Z Travel Limit (inches): 0.507 |
| --- | --- |
| Actuator Serial #: L157 Operator Name: Gordon, Matt Date: 4/13/2010 Time: 6:24 PM Actuator Coil Resistance: 6.45 Ohms, PASS Ambient Temperature: 73.3 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.525 Y Travel Limit (inches): 0.205 Z Travel Limit (inches): 0.503 | Actuator Serial #: L083 Operator Name: Gordon, Matt Date: 11/21/2009 Time: 5:52 PM Actuator Coil Resistance: 6.27 Ohms, PASS Ambient Temperature: 69.7 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.527 Y Travel Limit (inches): 0.205 Z Travel Limit (inches): 0.508 |
| Actuator Serial #: L144 Operator Name: Gordon, Matt Date: 4/13/2010 Time: 9:52 AM Actuator Coil Resistance: 6.39 Ohms, PASS Ambient Temperature: 73.31 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.526 Y Travel Limit (inches): 0.205 Z Travel Limit (inches): 0.508 | Actuator Serial #: L130 Operator Name: Gordon, Matt Date: 4/12/2010 Time: 1:52 PM Actuator Coil Resistance: 6.41 Ohms, PASS Ambient Temperature: 73.3 F Hi Pot Test Results: 1000 MOhms, PASS X Travel Limit (inches): 0.540 Y Travel Limit (inches): 0.208 Z Travel Limit (inches): 0.506 |

Acceptance Criteria:

- Actuators were previously tested and results are reported in T900564.

The tests report must contain:

Test results (Passed:    X   	Failed:        ).””

# Tests to be performed during assembly

## Step 1: Inventory (E1000052)

| DCC/Vendor number | Part name | Configuration | S/N | S/N | S/N |
| --- | --- | --- | --- | --- | --- |
| D071001 | Stage 0 base |  | 013 |  |  |
| D071051 | Stage 1 base |  | 014 |  |  |
| D071050 | Optical table |  | 014 |  |  |
| D071002 | Spring Post |  | 029 | 045 | 003 |
| D071100 | Spring |  | 037 | 043 | 048 |
| D071102 | Flexure |  | 003 | 012 | 020 |
| ADE | Position sensor | Horizontal | 11982 | 11995 | 12010 |
|  |  | Vertical | 12003 | 11993 | 12037 |
| D047812 | GS-13 pod | Horizontal | 29 | 85 | 90 |
|  |  | Vertical | 25 | 33 | 51 |
| D047823 | L4C pod | Horizontal |  |  |  |
|  |  | Vertical |  |  |  |
| D0902749 | Actuator | Horizontal | L133 | L157 | L144 |
|  |  | Vertical | L069 | L083 | L130 |

## Step 2: Check torques on all bolts

Acceptance Criteria:

- All bolts should trip the wrench, and start moving immediately after. If any bolts in a pattern move before torque is reached, recheck after all bolts are brought to spec.

Test result:	Passed:   X    	Failed:         .

## Step 3: Check gaps under Support Posts

Acceptance Criteria:

- A 0.001 inch shim cannot be passed freely through any connection to Stage 0 or between post and gussets. If shim can pass through, loosen all constraining bolts, and then retighten iteratively from the center of the part to the edges. Retest.

Test result:	Passed:   X   	Failed:         .

## Step 4: Pitchfork/Boxwork flatness before Optical Table install

Acceptance Criteria:

- Shim inserted won’t pass between parts.

Test result:	Passed:   X   	Failed:         .

## Step 5: Blade spring profile

| Blade # | Base (") | Tip(") | Flatness (mils) |
| --- | --- | --- | --- |
| 1 | 0.487 | 0.485 | + 2 |
| 2 | 0.503 | 0.494 | + 9 |
| 3 | 0.496 | 0.4875 | + 8.5 |

Table 1 - Blade profile

Acceptance Criteria:

- Blades must be flat within 0.015" inches.

Test result:	Passed:   X   	Failed:         .

## Step 6: Gap checks on actuators-after installation on Stage 1

| Actuator | Front Gap (1/1000") | Back Gap (1/1000") |
| --- | --- | --- |
| H1 | 0.085 | 0.080 |
| H2 | 0.085 | 0.085 |
| H3 | 0.080 | 0.080 |
| V1 | 0.085 | 0.085 |
| V2 | 0.080 | 0.080 |
| V3 | 0.085 | 0.085 |

Acceptance Criteria

- Gaps must be within 0.010” of design (i.e. 0.090” and .070” pass, but 0.095” and 0.065” doesn’t).

Test result:	Passed:   X   	Failed:         .

## Step 7: Check level of Stage 0

Max angle = (0.05)/72=

Note: These results were the ones before we leveled the test stand. We didn’t record the results after leveling it (but it was within specs). We will try to re-shoot Stage 0 and the Optical Table and record the numbers this time.

Acceptance Criteria

- The maximum angle of the table with the horizontal mustn’t exceed ~100µrad

Test result:	Passed:       	Failed:        . 	Waived:  X

## Step 8: Check level of Stage 1 Optical Table

Optical Level measurement of Stage 1 at large (12 - 15) number of points.

Max angle = (0.013)/85.59= 152 urad

Note: These results were the ones before we leveled the test stand. We didn’t record the results after leveling it (but it was within specs). We will try to re-shoot Stage 0 and the Optical Table and record the numbers this time.

Acceptance Criteria

- The maximum angle of the table with the horizontal mustn’t exceed ~100µrad

Test result:	Passed:      	Failed:          Waived:   X    >

## Step 9: Mass budget

Figure 7: Keel Masses and Wall masses location

|  |  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| w9 |  | 1 |  |  | 1 |  | 1 | 1 | 47.9 | 21.73 |
| w1 |  |  |  |  |  |  |  | 2 | 54.4 | 24.68 |
| w2 |  |  | 1 |  |  | 1 | 1 | 1 | 51.8 | 23.50 |
| w3 |  |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w4 |  |  |  |  |  |  | 2 | 1 | 58.4 | 26.49 |
| w5 |  |  |  | 1 |  |  |  | 2 | 56.6 | 25.67 |
| w6 |  | 1 |  |  |  |  | 1 | 1 | 43.4 | 19.69 |
| w7 |  |  | 1 | 1 |  |  |  | 2 | 57.7 | 26.17 |
| w8 |  |  |  |  |  |  | 2 | 1 | 58.4 | 26.49 |
| Side Masses Total |  | 2 | 2 | 2 | 1 | 1 | 8 | 12 | 471.4 | 213.82 |

Table 2: Wall masses distribution

|  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| k1 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k2 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
| k3 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k4 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
| k5 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k6 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
|  | 3 | 3 | 3 | 0 | 3 | 0 | 6 | 198.6 | 90.08 |

|  | Side | Keel | Top | Total |
| --- | --- | --- | --- | --- |
| Weigh (kg) | 213.82 | 90.08 | 277 | 580.9 |
| Torque x at O (N.m) |  | 0.00 | -69.02 |  |
| Torque y at O (N.m) |  | 5.26 | 0 |  |

Table 3: Masses distribution (computed using T1100261)

Acceptance Criteria

The Mass budget must be

- 579.1 Kg (cf E1100427)+/-25Kg (5%)

Test result:	Passed:   X   	Failed:         .

## Step 10: Shim thickness

| Lockers | Shim thickness (mils) |
| --- | --- |
| A | 129 |
| B | 130 |
| C | 127 |
| D | 125 |

Table 4: Shims Thickness

Acceptance Criteria

- Inventory is complete

Test result:	Passed:   X   	Failed:         .

## Step 11: Lockers adjustment

| D.I at Lockers | Dial indicators V | Dial indicators H |
| --- | --- | --- |
| A | -0.5 | -1 |
| B | 0 | 0 |
| C | -1 | 0 |
| D | 1 | -1.25 |

Table 5: Dial indicators read-out (in thousands of an inch)

Issues/difficulties encountered during this test : N/A

Acceptance Criteria

- Vertical and horizontal displacement near the lockers must be lower than 2 mils (0.002”)

Test result:	Passed:   X    	Failed:         .

# Tests to be performed after assembly

## Step 1 - Electronics Inventory

| Hardware | LIGO reference | S/N |
| --- | --- | --- |
| Coil driver | D0902744 | S1103354 |
|  |  | S1000316 |
| Anti Image filter | D070081 | S1000249 |
| Anti aliasing filter | D1000269 | S1106137 |
|  |  | S1106138 |
| Interface chassis | D1000067 | S1102219 |
|  |  | S1106356 |
|  |  | S1106358 |

Table 6: Inventory electronics

Acceptance Criteria

- Inventory is complete

Test result:	Passed:   X   	Failed:         .

## Step 2 - Set up sensors gap

|  | 10 Kg masses at each corners | No mass | No mass |
| --- | --- | --- | --- |
| Table locked | ADE boxes on | ADE boxes on | ADE boxes off |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation |
| H1 | 587.62 | 18.728 | 436.5 | 35.03 | 482.56 | 37.98 |
| H2 | -239.97 | 11.364 | -246.13 | 20.126 | -209.3 | 24.27 |
| H3 | 109.61 | 14.577 | 109.6 | 17.622 |  |  |
| V1 | 129.91 | 25.659 | 228.61 | 33.354 | 263.14 | 37.30 |
| V2 | 957.49 | 34.115 | 1019 | 43.678 | 1063 | 50.19 |
| V3 | 553.89 | 51.034 | 759.89 | 22.382 |  |  |

Capacitive position sensor readout after gap set-up

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- All mean values must be lower than 400 cts (a bit less than .0005”).

- All standard deviations below 5 counts.

- No cross talk

Test result:	Passed:     	Failed:    X    .

## Step 3 - Measure the Sensor gap

| Sensors | Gap measured on the Jig | Gap measured on the table |
| --- | --- | --- |
| H1 | NR | NR |
| H2 | NR | NR |
| H3 | NR | NR |
| V1 | NR | NR |
| V2 | NR | NR |
| V3 | NR | NR |

Acceptance criteria:

Sensors gap measured on the jig and on the optic table must be:

- 0.080” +/-0.002”

Test result:	Passed:      	Failed:         .

## Step 4 - Check Sensor gaps after the platform release

|  | Table locked | Table unlocked |
| --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Difference |
| H1 | 193.64 | 25.681 | 1151 | 957.36 |
| H2 | -236.13 | 26.746 | 154.17 | 390.3 |
| H3 | 151.94 | 25.349 | -201.41 | -353.35 |
| V1 | 163.02 | 17.694 | -103.31 | -266.33 |
| V2 | 28.829 | 35.983 | 1630.9 | 1602.071 |
| `V3 | 170.9 | 26.131 | 918.5 | 747.6 |

Table 7: Sensor gaps after platform release

Acceptance criteria:

- Absolute values of the difference between the unlocked and the locked table must be below:

  - 1600 cts for horizontal sensors (~0.002”)

  - 1600 cts for vertical sensors (~0.002”)

- Considering the acceptance criteria of step 4, all mean values must be lower than

  - 2000 cts for horizontal sensors (~0.0025”)

  - 2000 cts for vertical sensors (~0.0025”)

Comments:

Test result:	Passed:   X   	Failed:         .

## Step 5 – Performance of the limiter

## Step 5.1 - Test Nº1 - Push “in the general coordinates”

|  | CPS read out | Calculated after calibration |
| --- | --- | --- |
| Sensors | UP (Counts) | Down (Counts) | UP (mil) | Down (mil) |
| V1 | 20024 | -18943 | 23.8 | -22.6 |
| V2 | 20666 | -17876 | 24.6 | -21.3 |
| V3 | 20925 | -19376 | 24.9 | -23.1 |
|  | CPS read out | Calculated after calibration |
| Sensors | CW(-RZ) | CCW (+RZ) | CW (mil) | CCW (mil) |
| H1 | 21643 | -20224 | 25.8 | -24.1 |
| H2 | 22291 | -22916 | 26.5 | -27.3 |
| H3 | 22194 | -19112 | 26.4 | -22.8 |

Table 8: Optic table range of motion

## Step 5.2 - Test Nº2 – Push “locally”

|  | Push in positive direction | Push in negative direction | Railing | Actuator Gap Check |
| --- | --- | --- | --- | --- |
| H1 | -23492 | 24456 |  | X |
| H2 | -23879 | 23441 |  | X |
| H3 | -24668 | 25008 |  | X |
| V1 | -18884 | 20369 |  | X |
| V2 | -24970 | 27250 |  | X |
| V3 | -22098 | 22561 |  | X |

Table 9: Optic table range of motion

Acceptance criteria:

- The vertical sensor readout must be positive when the optic table is pushed in the +Z direction

- The horizontal sensor readout must be negative when the optic table is pushed in the +RZ direction

- Step 5.1

  - Absolutes value of all estimated motions must be higher than 16000counts (~0.020”)

- Step 5.2

  - No contact point on sensors

  - Absolute value of sensor read out must be higher than 16000counts (~0.020”)

  - No contact point on actuators

Test result:	Passed:   X   	Failed:         .

## Step 6 - Position Sensors unlocked/locked Power Spectrum

We swapped V3 CPS 12037 with 12058 and retook this test.

Data files in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/ Powerspectra/Undamped

- LLO_HAM_ISI_India_1_Calibrated_PSD_CPS_GS13_Unlocked_Locked_2013_12_23.mat

Scripts files for processing and plotting in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Scripts/Data_collection/

- Powerspectra_Measurements_Locked_Unlocked_HAM_ISI.m

- Powerspectra_Measurements_Tilted_HAM_ISI.m

Figures in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Figures/Powerspectra/Undamped

- LLO_HAM_ISI_India_1_Calibrated_PSD_CPS_Unlocked_Locked_2013_12_23.fig

- LLO_HAM_ISI_India_1_Calibrated_PSD_GS13_Unlocked_Locked_2013_12_23.fig

CPS calibration:

The CPS power spectrums are calibrated by using a sensitivity of 30.2 nm/count.

Figure 8: Calibrated CPS power spectrum

Figure 9: Power spectrum Calibrated GS13

Acceptance criteria:

- No cross talk (peaks at low frequencies + harmonics on measurements)

- Magnitudes of power spectra must be between requirement curves such as in the following figures (dashed lines)

| Sensors | ISI state | Frequency (Hz) | 2x10-2 | 1x10-1 | 1 | 10 | 20 | 100 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| GS-13 | Table locked | Max | 3x10-1 | 3x10-4 | 3x10-7 | 10-7 |  | 10-11 | 10-14 |
|  |  | Min | 3x10-4 | 3x10-7 | 3x10-10 | 10-12 |  | 10-14 | 10-17 |
|  | Table unlocked | Max | 1 | 3x10-3 | 10-5 | 10-9 |  | 10-11 | 10-14 |
|  |  | Min | 10-4 | 3x10-7 | 10-9 | 10-13 |  | 10-15 | 10-18 |
| CPS | Table locked | Max | 2*10-7 | 2x10-8 | 10-8 | 5x10-8 | 2x10-7 | 5x10-9 | 10-9 |
|  |  | Min | 5x10-9 | 2x10-9 | 8x10-10 | 5x10-10 |  | 10-10 | 5x10-11 |
|  | Table unlocked | Max | 2x10-6 | 8x10-7 | 8x10-7 | 5x10-8 | 2x10-7 | 2x10-8 | 10-9 |
|  |  | Min | 10-7 | 5x10-8 | 8x10-9 | 5x10-10 |  | 2x10-10 | 10-10 |

Table 10: Step 6 -Normal conditions-Sensors power spectra requirements

Test result:	Passed:    X   	Failed:          .

## Step 7 - GS13 power spectrum -tabled tilted

The figure below presents the GS13 power spectrum when the table is unlocked and loaded with a 20Kg mass at one of its corner.

Figure 10: Power spectrum Calibrated GS13 with mass at corner

Data files in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/ Powerspectra/Undamped

- LLO_HAM_ISI_India_1_Calibrated_PSD_GS13_Table_Tilted_2013_10_10.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Scripts/Data_collection/

- Powerspectra_Measurements_Tilted_HAM_ISI.m

Figures in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Figures/ Powerspectra/Undamped

- LLO_HAM_ISI_India_1_Calibrated_PSD_GS13_Table_Tilted_2013_10_10.fig

Acceptance criteria:

- With table unlocked and tilted, magnitudes of power spectra must be fully included within:

| Sensor | ISI State | Frequency | 5x10-1 Hz | 1 | 10 | 100 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| GS-13 | Table Tilted | Max | 2x10--5 | 2x10--6 | 8x10-8 | 4x10-11 | 3x10-14 |
|  |  | Min | 10-8 | 10-9 | 2x10-11 | 10-14 | 10-17 |

Table 11: Table Tilted- Sensors power spectra requirements

Test result:	Passed:      	Failed:    X     .

## Step 8- GS13 pressure readout

Scripts files for taking and processing the data, and plotting it in SVN at: seismic/SeiSVN/Common/MatlabTools

- gs13Presstest.m

Note: the figure hasn’t been plotted because of a bug in the script but the pressure sensors are all working fine and the pressure readouts are within requirements.

Issues/difficulties/comments regarding this test

Acceptance criteria:

- The pressure on all channels must be 25000 counts +/- 3000 counts

- All channels must follow comparable trend

Test result:	Passed:   X   	Failed:          .

## Step 9 - Coil Driver, cabling and resistance check

| Actuator | V1 | H1 | V2 |
| --- | --- | --- | --- |
| Coil driver | S1103354 – Fine 1 | S1103354 – Coarse 1 | S1103354 – Coarse 2 |
| Anti image pin # |  |  |  |
| Cable # | 28 | 25 | 30 |
| Resistance (Ohm) | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 |
|  | O.L (infinity) | 6.5 | O.L (infinity) | 6.6 | O.L (infinity) | 6.6 |
| MEDM offset (1000 counts) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) |
|  | 0.310 | 0.309 | 0.309 |
| Actuator | H2 | V3 | H3 |
| Coil driver | S1000317 - Coarse 2 | S1000316 - Coarse 1 | S1000316 - Fine 1 |
| Anti image pin # |  |  |  |
| Cable # | 29 |  |  |
| Resistance (Ohm) | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 | P1 - P2 | P2 - P3 |
|  | O.L (infinity) | 6.6 | O.L (infinity) | 6.6 | O.L (infinity) | 6.6 |
| MEDM offset (1000 counts) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) | Measurement P2 (-) ; P1&P3 (+) |
|  | 0.307 | 0.314 | 0.313 |

Table 12: Actuators resistance check

Acceptance criteria:

- The measured resistance between the middle pin and one side pin must be 6.5 +/-1 ohms

- Actuator neutral pins must be connected on pin #1 (left side pin of the plug)

- Actuator drive pins must be connected on pin #2 (middle pin of the plug)

- Actuator ground shield pins must be connected on pin #3 (right pin of the plug)

- All LEDs on the coil driver front panel must be green

The tests report must contain:

- The table “Actuators resistance check”

- Issues/difficulties/comments regarding this test

- Test result (Passed:    X   	Failed:          ).

## Step 10 - Actuators Sign and range of motion (Local drive)

|  | Negative drive | Positive drive |
| --- | --- | --- |
| H1 readout (count) | -23492 | 24456 |
| H2 readout (count) | -23879 | 23441 |
| H3 readout (count) | -24668 | 25008 |
| V1 readout (count) | -18884 | 20369 |
| V2 readout (count) | -24970 | 27250 |
| V3 readout (count) | -22098 | 22561 |

Table 13: Range of motion - Local drive

Data files in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Static_Tests

- LLO_HAM-ISI_India_1_Range_Of_Motion_1016201320131016.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Scripts/Data_collection

- Range_Motion_HAM_ISI.m

Acceptance criteria:

- Main couplings sensors readout must be at least 16000 counts (~0.02”)

- A positive offset drive on one actuator must give positive sensor readout on the collocated sensor. Signs will also be tested when measuring local to local transfer functions.

Test result:	Passed:   X   	Failed:         .

## Step 11 - Vertical Sensor Calibration

| Lockers | D.I readout with for a negative drive | D.I readout without any drive | D.I readout with for a positive drive |  |
| --- | --- | --- | --- | --- |
| A | 20 | 0 | -19 |  |
| B | 21 | 0 | -20 |  |
| C | 18 | 0 | -19 |  |
| D | 18 | 0 | -19 |  |
| Average | 19.25 | 0 | -19.25 | 38.5 |
| Sensors | Counts | Counts | Counts | Difference (Counts) |
| V1 | -16108 | -459.64 | 16397 | 32505 |
| V2 | -15132 | 1218.9 | 17635 | 32767 |
| V3 | -16104 | 390.69 | 15764 | 31868 |
|  | -15781.33 | 383.32 | Average | 32380.00 |

Table 14: Calibration of capacitive position sensors

Vertical sensitivity:	32380/38.5 = 841 count/mil

or	841 count/mil / 1638 V/count = 0.513V/mil

or	25400nm/mil / 841 mil/count = 30.20 nm/count

Nominal Calibration

CPS Sensitivity: 	20V/0.039" = 20V/39mils = 0.513V/mil

Calibration in counts: 	215 / 20 * 20/39 = 840 count/mil

or	25400 nm/mil * 1/840 mil/count = 30.2nm/count

Difference with Nominal sensitivity = (840-841)/840=-0.12%

Acceptance criteria:

- Deviation from nominal value < 2%. Nominal value is 840 count/mil.

Test result:	Passed:   X   	Failed:          .

## Step 12 - Vertical Spring Constant

Results presented below are obtained after the initial sensors calibration.

| Sensors | Initial position | Measurement 1 | Measurement 2 | Measurement 3 | Mean diff counts | Mean diff m | K (N/m) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| V1 | -783.96 | -8618.7 | -8550.5 | -8544.7 | -7787.34 | -2.35E-04 | 8.34E+04 |
| V2 | 790.7 | -7009.4 | -7047.6 | -7053.3 | -7827.47 | -2.36E-04 | 8.30E+04 |
| V3 | 273.08 | -7911.7 | -7905.9 | -7899 | -8178.61 | -2.47E-04 | 7.94E+04 |
|  |  |  |  |  |  | Average (N/m) | 2.46E+05 |

Table 15: Vertical spring constant

Acceptance criteria:

- +/-2 % of 2.4704e5 N/m (i.e. between 2.421e5 and 2.520e5 N/m)

- +/- 5% of variation between each spring and the average

The measured error on the vertical stiffness is 1.47%

Test result:	Passed:   X   	Failed:

## Step 13 - Static Testing (Tests in the local basis)

|  |  | H1 | H2 | H3 | V1 | V2 | V3 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Actuators (1000 counts) | H1 | -2121.57 | -1321.94 | -1267.80 | 103.79 | 85.01 | -18.71 |
|  | H2 | -1340.91 | -2063.54 | -1322.60 | 47.19 | -4.63 | 30.95 |
|  | H3 | -1368.23 | -1303.04 | -2084.10 | 49.69 | 14.17 | 7.25 |
|  | V1 | -239.44 | -129.64 | 324.40 | -1558.91 | -36.13 | 671.36 |
|  | V2 | 311.89 | -44.67 | -354.28 | 594.29 | -1579.93 | 15.16 |
|  | V3 | -276.18 | 465.20 | -209.53 | 90.59 | 726.53 | -1526.31 |

Table 16: Main and cross coupling

Data files in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Static_Tests

- LLO_HAM_ISI_India_1_Sensor_Readout_Local_20111126.mat

Scripts files for taking data in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Scripts/Data_Collection

- Sensor_Readout_Local_Drive_MEDM_HAM_ISI.m

Note: the compensation filters were not on while taking this test because they were tripping the watchdogs. This explains why all the results in the table have been multiplied by (-1).

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Vertical

For a +1000 count offset drive on vertical actuators

  - Collocated sensors must be 1400 counts +/- 10%

- Horizontal

For a +1000 count offset drive on horizontal actuators

  - Collocated sensors must be 2000 counts +/- 10%

  - Non-collocated horizontal sensors must be 1250 counts +/-10%

Test result:	Passed:       	Failed:    X     .

## Step 14 - Linearity test

|  | Slope | Offset | Average slope | Variation from average (%) |
| --- | --- | --- | --- | --- |
| H1 | 2.094 | 1963 | 2.0873 | 0.32 |
| H2 | 2.064 | 813 |  | -1.12 |
| H3 | 2.104 | 383 |  | 0.80 |
| V1 | 1.514 | -732 | 1.5050 | 0.60 |
| V2 | 1.499 | 826 |  | -0.40 |
| V3 | 1.502 | 161 |  | -0.20 |

Table 17: Slopes and offset of the triplet Actuators - HAM-ISI - Sensors

Scripts files for taking data in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Scripts/Data_Collection

- Linearity_Test_Awgstream_HAM_ISI.m

Figure 11: Horizontal and vertical actuators x HAM-ISI x sensors

Data files in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Linearity_test

- LLO_HAM_ISI_India_1_Linearity_test_20131016.mat

Figures in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Figures/Linearity_test

- LLO_HAM_ISI_India_1_Linearity_test_20131016.fig

- LLO_HAM_ISI_India_1_Linearity_test_20131016.pdf

Acceptance criteria:

- Horizontal and vertical slopes of the triplet actuators x HAM-ISI x sensors: Average slope +/- 2%

Test result:	Passed:   X   	Failed:         .

## Step 15 - Cartesian Basis Static Testing

|  | 1000 counts drive | X Drive | Y Drive | Z Drive | Rx Drive | Ry Drive | Rz Drive |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Sensors readout (count) | H1 | -245.489 | 509.831 | 48.381 | 564.031 | 357.551 | 1935.331 |
|  | H2 | -211.228 | -340.528 | 55.6723 | -417.728 | 423.5823 | 2131.572 |
|  | H3 | 539.79 | 20.41 | -67.07 | 59.49 | -626.37 | 1946.03 |
|  | V1 | -22.3 | 14.7 | -203.3 | 529.7 | 1730.29 | -54.3 |
|  | V2 | 12.9 | 48.6 | -352.1 | -1820.1 | -398.1 | -20.1 |
|  | V3 | -32.79 | 4.31 | -350.79 | 1261.41 | -1307.39 | 17.81 |
|  | Direction read out | -15807.1 | -15804 | -8314 | -81975 | -81878 | -78984 |

Table 18: Tests in the general coordinate basis

Issues/difficulties/comments regarding this test: the compensation filters were not on while taking this test because they were tripping the watchdogs. This explains why all the results in the table have been multiplied by (-1).

The results in the table are very consistent with what we had for the previous units. The only issue we have on this test is the direction read out which is with the direction read out values, the calibration filters were on so these values are in nm instead of counts, explaining the high numbers.

Acceptance criteria:

For a positive drive in the Cartesian basis:

- Local sensor readout must have the same sign that the reference table (CONT2ACT check)

- Cartesian sensors read out must be positive (DISP2CEN check) in the drive direction

Test result:	Passed:    X   	Failed:         .

## Step 16- Frequency response

Compensation filters of the new GS13 interface chassis are located in the geophone pre-filters bank. Powerspectra were measured with masses on the optic table not bolted.

## Step 16.1 - Local to local measurements

Local to local transfer functions have been measured with 90 repetitions.

Data files in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Transfer_Functions/Measurements/

Undamped/

- LLO_HAM_ISI_India_1_Data_TF_L2L_50mHz_500mHz_20131125-191024.mat

- LLO_HAM_ISI_India_1_Data_TF_L2L_500mHz_5Hz_20131125-160450.mat

- LLO_HAM_ISI_India_1_Data_TF_L2L_5Hz_200Hz_20131121-144422.mat

- LLO_HAM_ISI_India_1_Data_TF_L2L_200Hz_800Hz_20131121-131047.mat

Data collection script files:

ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Scripts/Data_Collection

- Run_TF_L2L_50mHz_500mHz.m

- Run_TF_L2L_500mHz_5Hz.m

- Run_TF_L2L_5mHz_200Hz.m

- Run_TF_L2L_200Hz_800Hz.m

Scripts files for processing and plotting in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Transfer_Functions/Measurements/

Undamped/

- Plot_LLO_HAM_ISI_India_1_TF_L2L_2013_10_25.m

Figures in SVN at:

ligo/svncommon/SeiSVN/seismic/HAMISI/X2/Data/India_1/Figures/Transfer_Functions/Measurements/Undamped/

- LLO_HAM_ISI_India_1_TF_L2L_V_CPS_50mHz_800Hz_2013_11_21.fig

- LLO_HAM_ISI_India_1_TF_L2L_H_GS13_50mHz_800Hz_2013_11_21.fig

- LLO_HAM_ISI_India_1_TF_L2L_H_CPS_50mHz_800Hz_2013_11_21.fig

- LLO_HAM_ISI_India_1_TF_L2L_V_GS13_50mHz_800Hz_2013_11_21.fig

Storage of measured transfer functions in the SVN at:

ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Transfer_functions/ Measurements/

Undamped/

- LLO_HAM_ISI_India_1_Data_TF_L2L_2013_11_21.mat

The local to local transfer functions are presented below.

Table 19: Local to Local Measurements – Horizontal capacitive sensors

Table 20: Local to Local Measurements – Horizontal inertial sensors

Table 21: Local to Local Measurements – Vertical capacitive sensors

Table 22: Local to Local Measurements – Vertical inertial sensors

Issues/difficulties/comments regarding this test:

Around 1 Hz, the inertial sensors seem to indicate different behaviors of each corner.

## Step 16.2 - Cartesian to Cartesian measurements

Issues/difficulties/comments regarding this test:

Acceptance criteria:

- Local to local measurements

  - On CPS, the phase must be 0º at DC

  - On Geophones, the phase must be -90º at DC

  - Identical shape in each corner

- Cartesian to Cartesian measurements

  - On CPS, the phase must be 0º at DC

  - On Geophones, the phase must be -90º at DC

  - Identical shape X/Y and RX/RY

Test result:	Passed:      	Failed:         .	Waived:  X

## Step 17 - Transfer function comparison with Reference

## Step 17.1 - Local to local - Comparison with Reference

This is the 2nd unit which is compared to LHO Unit #2 instead of LLO HAM 6 (v4 was comparing to LLO HAM 6 and both data can be found on the SVN).

Scripts files for processing and plotting in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Transfer_Functions/Measurements/

Undamped/

- Plot_LLO_HAM_ISI_India_1_TF_L2L_2011_08_02

Local to local figures in SVN at:

/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Figures/Transfer_Functions/Measurements/Undamped

- LLO_HAM_ISI_India_1_TF_L2L_H_CPS_50mHz_800Hz_wRef_2011_08_02.fig

- LLO_HAM_ISI_India_1_TF_L2L_V_CPS_50mHz_800Hz_wRef_2011_08_02.fig

- LLO_HAM_ISI_India_1_TF_L2L_H_GS13_50mHz_800Hz_wRef_2011_08_02.fig

- LLO_HAM_ISI_India_1_TF_L2L_V_GS13_50mHz_800Hz_wRef_2011_08_02.fig

GS13, Local to local measurement

Figure - Local to local measurements comparison with LHO UNIT 2 – Horizontal GS-13

Figure - Local to local measurements comparison with LHO UNIT 2 – Vertical GS13

CPS, Local to local measurement, Undamped

Figure - Local to local measurements comparison – Horizontal Position sensors

Figure - Local to local measurements comparison – Vertical Position sensors

## Step 17.2 - Cartesian to Cartesian - Comparison with Reference

Scripts files for processing and plotting in SVN at:

ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Transfer_Functions/Measurements/

Undamped/

- Plot_LLO_HAM_ISI_India_1_TF_C2C_2011_08_02.m

Cartesian to Cartesian figures in SVN at :

/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Figures/Transfer_functions/Measurements/Undamped

Acceptance criteria:

- No difference with the reference transfer functions (SVN)

  - Phase – less than 10º - In Phase – Out of Phase

  - Damping (fit by eye with Reference transfer functions)

  - DC gain

  - Eigen frequencies shift less than 10%

Test result:	Passed:      	Failed:         .

## Step 18 - Lower Zero Moment Plane

Scripts files for processing and plotting in SVN at:

ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2Misc

- TF_2_LZMP.m

Figures in SVN at:

ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Misc/

- LLO_India_Unit_1_LZMP_26-Oct-2014.fig

- LLO_India_Unit_1_LZMP_26-Oct-2014.pdf

LMZP Offset X - mm

1.2815

LMZP Offset Y - mm

1.6562

Issues/difficulties/comments regarding this test:

LZMP data was missing. It was processed using the low frequency part of the existing L2L Transfer functions, applying the input/output filters to that, and pushing it into the Cartesian basis before doing the regular LZMP computation

LZMP plotting functions displays phase and coherence in a non-ideal way. This is very minor, and not worth updating for now.

Acceptance criteria:

- X offset must be less than 2 mm

- Y offset must be less than 2 mm

Test result:	Passed:   X    	Failed:         .

## Step 19 - Damping loops

In this step, HAM6 damping loops are implemented. First, damping performances are evaluated in simulation. Second, Damping loops are implemented and performance is experimentally measured.

## Step 19.1 - Transfer functions - Simulation

Scripts files for processing and plotting in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Transfer_Functions/Simulation/Damping

Figures in SVN at:

/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Figures/Transfer_Functions/Simulations/

Damping/

Results are saved in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Transfer_Functions/Simulations/

Damping/

Acceptance criteria:

- HAM6 damping loops must implemented and stable with

  - Phase margin must be at least 45º

  - Gain margin must be at least 20dB

Test result:	Passed:      	Failed:         .

## Step 19.2 - Powerspectra – Experimental

Data files in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Powerspectra/Damping/

Scripts files for taking data and plotting in SVN at:

/ligo/svncommon/SeiSVN/seismic/HAM-ISI/X2/Scripts/Data_Collection/

- Powerspectra_Measurements_Undamped_Damped_HAM_ISI.m

Figures in SVN at:

/svncommon/SeiSVN/seismic/HAM-ISI/X2/Data/India_1/Figures/Powerspectra/Damping/

Acceptance criteria:

- HAM6 damping loop must stable when all damping loops are engaged

- Similar damping effect than in simulated plots

Test result:	Passed:   X   	Failed:         .

Conclusion

A few issues were found during the testing of this HAM India 1. The known issues are summed up here:

- Sensor gaps not recorded on the jig (III Step 2 & 3)

- Level of Stage 0 and Optical Table (II Step 7 & 8) are within range but the numbers haven’t been recorded. It will have to be done.

- The Cartesian Basis Static Testing (III Step 15) direction readout is in nm instead of counts (issue with the script that needs to be fixed).

- The Static Testing Local Basis & Cartesian Basis (III Step 13 & 15) has been done with the compensation filters off because they were tripping the watchdogs. This explains why all the results in the table have been multiplied by (-1).

When this Unit is approved, it will be put in a shipping container and stored to be shipped to India later.

# Text boxes

-0.04”

-0.05”

0

Locker C

Locker C

-0.010”

-0.010”

0.003”

0

0

N

-0.003”