## Outgassing Rates
A-OSEM Head with Wound Coil (404 units)
Species | AMU | Outgassing Rate (Torr-L/s per AOSEM body)
H | 2 | 3.6702e-12
He | 4 | 1.2617e-15
C+ | 12 | 9.6072e-15
N+, N2+ | 14 | 2.2171e-14
water (O+) | 16 | 1.8208e-14
water (HO+) | 17 | 8.6091e-15
water (H2O+) | 18 | 2.951e-14
CO2++ | 22 | 9.1939e-16
N2 | 28 | 3.707e-13
O2 | 32 | 9.8456e-16
Ar | 40 | 3.3077e-15
hydrocarbon | 41 | 1.0013e-15
hydrocarbon | 43 | 9.9542e-16
CO2 | 44 | 9.6659e-15
hydrocarbon | 53 | 9.4111e-16
hydrocarbon | 55 | 8.4104e-16
hydrocarbon | 57 | 8.4797e-16
Sum HC |  | 4.6268e-15
| AMU |  | Outgassing Rate (Torr-L/s) per AOSEM Wound Head 404 units
hydrogen (H2) | 2 |  | 3.6702e-12
water (H2O) at 1000 hr | max of 16, 17, 18 |  | 2.951e-14
air (N2) | 28 |  | 3.707e-13
CO2 | 44 |  | 9.6659e-15
hydrocarbons | sum of 41, 43, 53, 55, 57 |  | 4.6268e-15

## F82310-ALIGO OSEMS
2010-08-23 | 2010 04:28:41 PM
Residual Gas Analyzer Software
Software Version | 3.206.001
Analog Scan Setup:
Data Points in Scan | 991
Units | Amps
Noise Floor | 1
CEM Status | OFF
Points Per AMU | 10
Scan Start Mass | 1 | amu
Scan Stop Mass | 100 | amu
Focus Voltage | 90 | Volts
Ion Energy | HIGH
Electron Energy | 70 | eV
CEM Voltage | 1000 | Volts
CEM Gain | 1000000
Sensitivity Factor | 0.0001
Filament Current | 1 | mAmps
-2.828e-17
1 | 1.62e-17 | 4.448e-17
1.1 | 7.59e-17 | 1.0418e-16
1.2 | 1.85e-16 | 2.1328e-16
1.3 | 1.03e-15 | 1.0583e-15
1.4 | 2.58e-15 | 2.6083e-15
1.5 | 1.22e-14 | 1.2228e-14
1.6 | 2.75e-14 | 2.7528e-14
1.7 | 6.8e-14 | 6.8028e-14
1.8 | 1.06e-13 | 1.0603e-13
1.9 | 1.19e-13 | 1.1903e-13
2 | 1.25e-13 | 1.2503e-13
2.1 | 1.2e-13 | 1.2003e-13
2.2 | 1.01e-13 | 1.0103e-13
2.3 | 7.03e-14 | 7.0328e-14
2.4 | 1.99e-14 | 1.9928e-14
2.5 | 1.21e-15 | 1.2383e-15
2.6 | 2.99e-16 | 3.2728e-16
2.7 | 1.25e-16 | 1.5328e-16
2.8 | 6.3e-17 | 9.128e-17
2.9 | 6.48e-18 | 3.476e-17
3 | -1.01e-17 | 1.818e-17
3.1 | 5.57e-18 | 3.385e-17
3.2 | 1.22e-17 | 4.048e-17
3.3 | 1.13e-17 | 3.958e-17
3.4 | 6.44e-18 | 3.472e-17
3.5 | 2.16e-18 | 3.044e-17
3.6 | -3.37e-20 | 2.8246e-17
3.7 | 6.75e-19 | 2.8955e-17
3.8 | 6.21e-18 | 3.449e-17
3.9 | 1.07e-17 | 3.898e-17
4 | 1.47e-17 | 4.298e-17
4.1 | 1.79e-17 | 4.618e-17
4.2 | 2.03e-17 | 4.858e-17
4.3 | 1.77e-17 | 4.598e-17
4.4 | 1.08e-17 | 3.908e-17
4.5 | 5.83e-18 | 3.411e-17
4.6 | 1.28e-18 | 2.956e-17
4.7 | 1.18e-18 | 2.946e-17
4.8 | 2.8e-18 | 3.108e-17
4.9 | 3e-18 | 3.128e-17
5 | 1.55e-18 | 2.983e-17
5.1 | 3.34e-18 | 3.162e-17
5.2 | 1.01e-18 | 2.929e-17
5.3 | 1.96e-18 | 3.024e-17
5.4 | 4.69e-18 | 3.297e-17
5.5 | 4.32e-18 | 3.26e-17
5.6 | 2.93e-18 | 3.121e-17
5.7 | 4.32e-18 | 3.26e-17
5.8 | 5.73e-18 | 3.401e-17
5.9 | 1.25e-18 | 2.953e-17
6 | 3.78e-18 | 3.206e-17
6.1 | 6.75e-19 | 2.8955e-17
6.2 | -1.89e-18 | 2.639e-17
6.3 | -3.04e-18 | 2.524e-17
6.4 | 2.39e-18 | 3.067e-17
6.5 | -7.76e-19 | 2.7504e-17
6.6 | 2.46e-18 | 3.074e-17
6.7 | 1.89e-18 | 3.017e-17
6.8 | 2.6e-18 | 3.088e-17
6.9 | 1.89e-18 | 3.017e-17
7 | 1.42e-18 | 2.97e-17
7.1 | 5.9e-18 | 3.418e-17
7.2 | 4.99e-18 | 3.327e-17
7.3 | 5.26e-18 | 3.354e-17
7.4 | 2.26e-18 | 3.054e-17
7.5 | 1.42e-18 | 2.97e-17
7.6 | 2.63e-18 | 3.091e-17
7.7 | 2.39e-18 | 3.067e-17
7.8 | 2.73e-18 | 3.101e-17
7.9 | 3.04e-18 | 3.132e-17
8 | 3.61e-18 | 3.189e-17
8.1 | 5.4e-19 | 2.882e-17
8.2 | 2.23e-18 | 3.051e-17
8.3 | 6.75e-20 | 2.8348e-17
8.4 | -3.1e-18 | 2.518e-17
8.5 | -1.35e-18 | 2.693e-17
8.6 | 8.43e-19 | 2.9123e-17
8.7 | 2.36e-18 | 3.064e-17
8.8 | 1.48e-18 | 2.976e-17
8.9 | 0 | 2.828e-17
9 | 2.5e-18 | 3.078e-17
9.1 | 3.98e-18 | 3.226e-17
9.2 | 5.3e-18 | 3.358e-17
9.3 | 5.8e-18 | 3.408e-17
9.4 | 4.52e-18 | 3.28e-17
9.5 | 3.2e-18 | 3.148e-17
9.6 | 6.14e-18 | 3.442e-17
9.7 | 4.65e-18 | 3.293e-17
9.8 | 9.78e-19 | 2.9258e-17
9.9 | 5.46e-18 | 3.374e-17
10 | 6.04e-18 | 3.432e-17
10.1 | 3.68e-18 | 3.196e-17
10.2 | 4.69e-18 | 3.297e-17
10.3 | 6.64e-18 | 3.492e-17
10.4 | 2.97e-18 | 3.125e-17
10.5 | -4.72e-19 | 2.7808e-17
10.6 | 4.72e-19 | 2.8752e-17
10.7 | 2.66e-18 | 3.094e-17
10.8 | 3e-18 | 3.128e-17
10.9 | 3.81e-18 | 3.209e-17
11 | 3.34e-18 | 3.162e-17
11.1 | 5.67e-18 | 3.395e-17
11.2 | 4.35e-18 | 3.263e-17
11.3 | 3.61e-18 | 3.189e-17
11.4 | 4.05e-18 | 3.233e-17
11.5 | 4.35e-18 | 3.263e-17
11.6 | 3.02e-17 | 5.848e-17
11.7 | 1.33e-16 | 1.6128e-16
11.8 | 1.92e-16 | 2.2028e-16
11.9 | 2.49e-16 | 2.7728e-16
12 | 2.99e-16 | 3.2728e-16
12.1 | 3.05e-16 | 3.3328e-16
12.2 | 2.81e-16 | 3.0928e-16
12.3 | 2.19e-16 | 2.4728e-16
12.4 | 1.14e-16 | 1.4228e-16
12.5 | 5.83e-17 | 8.658e-17
12.6 | 2.5e-18 | 3.078e-17
12.7 | -1.84e-17 | 9.88e-18
12.8 | 6.04e-18 | 3.432e-17
12.9 | 6.91e-17 | 9.738e-17
13 | 8.19e-17 | 1.1018e-16
13.1 | 6.79e-17 | 9.618e-17
13.2 | 5.98e-17 | 8.808e-17
13.3 | 5.74e-17 | 8.568e-17
13.4 | 3.57e-17 | 6.398e-17
13.5 | -2.51e-17 | 3.18e-18
13.6 | 4.35e-17 | 7.178e-17
13.7 | 2.14e-16 | 2.4228e-16
13.8 | 4.41e-16 | 4.6928e-16
13.9 | 6.36e-16 | 6.6428e-16
14 | 7.27e-16 | 7.5528e-16
14.1 | 7.38e-16 | 7.6628e-16
14.2 | 6.69e-16 | 6.9728e-16
14.3 | 4.98e-16 | 5.2628e-16
14.4 | 2.42e-16 | 2.7028e-16
14.5 | 1.06e-16 | 1.3428e-16
14.6 | 6.84e-17 | 9.668e-17
14.7 | 9.04e-17 | 1.1868e-16
14.8 | 1.96e-16 | 2.2428e-16
14.9 | 3.14e-16 | 3.4228e-16
15 | 3.95e-16 | 4.2328e-16
15.1 | 4.2e-16 | 4.4828e-16
15.2 | 3.82e-16 | 4.1028e-16
15.3 | 3.13e-16 | 3.4128e-16
15.4 | 1.89e-16 | 2.1728e-16
15.5 | 1.01e-16 | 1.2928e-16
15.6 | 7.25e-17 | 1.0078e-16
15.7 | 1.3e-16 | 1.5828e-16
15.8 | 3.09e-16 | 3.3728e-16
15.9 | 4.83e-16 | 5.1128e-16
16 | 5.92e-16 | 6.2028e-16
16.1 | 6.13e-16 | 6.4128e-16
16.2 | 5.81e-16 | 6.0928e-16
16.3 | 4.6e-16 | 4.8828e-16
16.4 | 2.58e-16 | 2.8628e-16
16.5 | 1.13e-16 | 1.4128e-16
16.6 | 6.62e-17 | 9.448e-17
16.7 | 8.35e-17 | 1.1178e-16
16.8 | 1.51e-16 | 1.7928e-16
16.9 | 2.27e-16 | 2.5528e-16
17 | 2.65e-16 | 2.9328e-16
17.1 | 2.79e-16 | 3.0728e-16
17.2 | 2.64e-16 | 2.9228e-16
17.3 | 2.17e-16 | 2.4528e-16
17.4 | 1.28e-16 | 1.5628e-16
17.5 | 7.64e-17 | 1.0468e-16
17.6 | 8e-17 | 1.0828e-16
17.7 | 2.12e-16 | 2.4028e-16
17.8 | 5.36e-16 | 5.6428e-16
17.9 | 8.41e-16 | 8.6928e-16
18 | 9.77e-16 | 1.0053e-15
18.1 | 9.72e-16 | 1.0003e-15
18.2 | 8.92e-16 | 9.2028e-16
18.3 | 6.8e-16 | 7.0828e-16
18.4 | 3.35e-16 | 3.6328e-16
18.5 | 1.36e-16 | 1.6428e-16
18.6 | 7.87e-17 | 1.0698e-16
18.7 | 7.18e-17 | 1.0008e-16
18.8 | 9.46e-17 | 1.2288e-16
18.9 | 1.13e-16 | 1.4128e-16
19 | 1.37e-16 | 1.6528e-16
19.1 | 1.53e-16 | 1.8128e-16
19.2 | 1.57e-16 | 1.8528e-16
19.3 | 1.28e-16 | 1.5628e-16
19.4 | 7.73e-17 | 1.0558e-16
19.5 | 2.27e-17 | 5.098e-17
19.6 | -2.6e-17 | 2.28e-18
19.7 | -1.32e-17 | 1.508e-17
19.8 | -2.87e-18 | 2.541e-17
19.9 | 5.26e-18 | 3.354e-17
20 | 8.3e-18 | 3.658e-17
20.1 | 1.01e-17 | 3.838e-17
20.2 | 1.22e-17 | 4.048e-17
20.3 | 9.88e-18 | 3.816e-17
20.4 | 4.52e-18 | 3.28e-17
20.5 | 2.83e-18 | 3.111e-17
20.6 | 5.4e-19 | 2.882e-17
20.7 | 3.37e-18 | 3.165e-17
20.8 | 6.91e-18 | 3.519e-17
20.9 | 3.71e-18 | 3.199e-17
21 | 2.02e-18 | 3.03e-17
21.1 | 2.83e-18 | 3.111e-17
21.2 | 1.69e-18 | 2.997e-17
21.3 | -3.37e-20 | 2.8246e-17
21.4 | 1.69e-18 | 2.997e-17
21.5 | 1.08e-18 | 2.936e-17
21.6 | 2.06e-18 | 3.034e-17
21.7 | 7.08e-19 | 2.8988e-17
21.8 | 3.91e-18 | 3.219e-17
21.9 | 3.95e-18 | 3.223e-17
22 | 3.04e-18 | 3.132e-17
22.1 | 2.63e-18 | 3.091e-17
22.2 | 3.17e-18 | 3.145e-17
22.3 | 4.05e-18 | 3.233e-17
22.4 | 2.73e-18 | 3.101e-17
22.5 | 6.37e-18 | 3.465e-17
22.6 | 7.32e-18 | 3.56e-17
22.7 | 1.1e-17 | 3.928e-17
22.8 | 1.09e-17 | 3.918e-17
22.9 | 9.98e-18 | 3.826e-17
23 | 1.22e-17 | 4.048e-17
23.1 | 1.27e-17 | 4.098e-17
23.2 | 9.68e-18 | 3.796e-17
23.3 | 1.06e-17 | 3.888e-17
23.4 | 7.02e-18 | 3.53e-17
23.5 | 3.41e-18 | 3.169e-17
23.6 | 2.6e-18 | 3.088e-17
23.7 | 2.46e-18 | 3.074e-17
23.8 | 3.31e-18 | 3.159e-17
23.9 | 6.54e-18 | 3.482e-17
24 | 7.72e-18 | 3.6e-17
24.1 | 5.23e-18 | 3.351e-17
24.2 | 5.03e-18 | 3.331e-17
24.3 | 3.24e-18 | 3.152e-17
24.4 | 3.34e-18 | 3.162e-17
24.5 | 4.55e-18 | 3.283e-17
24.6 | 3.37e-18 | 3.165e-17
24.7 | 7.89e-18 | 3.617e-17
24.8 | 1.45e-17 | 4.278e-17
24.9 | 1.47e-17 | 4.298e-17
25 | 1.09e-17 | 3.918e-17
25.1 | 1.55e-17 | 4.378e-17
25.2 | 7.41e-17 | 1.0238e-16
25.3 | 5.74e-17 | 8.568e-17
25.4 | -5.23e-18 | 2.305e-17
25.5 | -2.02e-17 | 8.08e-18
25.6 | -6.64e-18 | 2.164e-17
25.7 | 6.68e-18 | 3.496e-17
25.8 | 8.86e-17 | 1.1688e-16
25.9 | 9.24e-17 | 1.2068e-16
26 | 9.04e-17 | 1.1868e-16
26.1 | 8.67e-17 | 1.1498e-16
26.2 | 8.31e-17 | 1.1138e-16
26.3 | 7.78e-17 | 1.0608e-16
26.4 | 5.99e-17 | 8.818e-17
26.5 | 7.72e-18 | 3.6e-17
26.6 | -2.11e-17 | 7.18e-18
26.7 | -2.7e-18 | 2.558e-17
26.8 | 5.48e-17 | 8.308e-17
26.9 | 1.02e-16 | 1.3028e-16
27 | 1.04e-16 | 1.3228e-16
27.1 | 1.21e-16 | 1.4928e-16
27.2 | 1.39e-16 | 1.6728e-16
27.3 | 1.89e-16 | 2.1728e-16
27.4 | 2.69e-16 | 2.9728e-16
27.5 | 4.46e-16 | 4.7428e-16
27.6 | 1.16e-15 | 1.1883e-15
27.7 | 5.13e-15 | 5.1583e-15
27.8 | 9.59e-15 | 9.6183e-15
27.9 | 1.17e-14 | 1.1728e-14
28 | 1.26e-14 | 1.2628e-14
28.1 | 1.2e-14 | 1.2028e-14
28.2 | 1.06e-14 | 1.0628e-14
28.3 | 7.31e-15 | 7.3383e-15
28.4 | 2.19e-15 | 2.2183e-15
28.5 | 3.92e-16 | 4.2028e-16
28.6 | 1.53e-16 | 1.8128e-16
28.7 | 9.19e-17 | 1.2018e-16
28.8 | 8.21e-17 | 1.1038e-16
28.9 | 9.57e-17 | 1.2398e-16
29 | 1.09e-16 | 1.3728e-16
29.1 | 1.1e-16 | 1.3828e-16
29.2 | 1.07e-16 | 1.3528e-16
29.3 | 9.81e-17 | 1.2638e-16
29.4 | 7.05e-17 | 9.878e-17
29.5 | 1.84e-17 | 4.668e-17
29.6 | -2.48e-17 | 3.48e-18
29.7 | -8.3e-18 | 1.998e-17
29.8 | 2.16e-18 | 3.044e-17
29.9 | 8.26e-18 | 3.654e-17
30 | 1.42e-17 | 4.248e-17
30.1 | 1.51e-17 | 4.338e-17
30.2 | 1.65e-17 | 4.478e-17
30.3 | 1.85e-17 | 4.678e-17
30.4 | 8.06e-17 | 1.0888e-16
30.5 | 2.86e-17 | 5.688e-17
30.6 | -2.74e-17 | 8.8e-19
30.7 | -1.55e-17 | 1.278e-17
30.8 | -3.88e-18 | 2.44e-17
30.9 | 2.36e-18 | 3.064e-17
31 | 2.33e-18 | 3.061e-17
31.1 | 2.36e-18 | 3.064e-17
31.2 | 1.32e-18 | 2.96e-17
31.3 | 1.69e-18 | 2.997e-17
31.4 | 1.11e-18 | 2.939e-17
31.5 | 1.18e-18 | 2.946e-17
31.6 | 4.01e-18 | 3.229e-17
31.7 | 3.98e-18 | 3.226e-17
31.8 | 5.5e-18 | 3.378e-17
31.9 | 4.01e-18 | 3.229e-17
32 | 5.26e-18 | 3.354e-17
32.1 | 5.5e-18 | 3.378e-17
32.2 | 4.01e-18 | 3.229e-17
32.3 | 8.43e-18 | 3.671e-17
32.4 | 7.93e-18 | 3.621e-17
32.5 | 5.3e-18 | 3.358e-17
32.6 | 4.52e-18 | 3.28e-17
32.7 | 6e-18 | 3.428e-17
32.8 | 1.32e-18 | 2.96e-17
32.9 | 4.01e-18 | 3.229e-17
33 | 5.97e-18 | 3.425e-17
33.1 | 3.51e-18 | 3.179e-17
33.2 | 4.15e-18 | 3.243e-17
33.3 | 6.81e-18 | 3.509e-17
33.4 | 9.31e-18 | 3.759e-17
33.5 | 6.44e-18 | 3.472e-17
33.6 | 2.06e-18 | 3.034e-17
33.7 | 1.75e-18 | 3.003e-17
33.8 | 2.7e-19 | 2.855e-17
33.9 | 3.41e-18 | 3.169e-17
34 | 6.95e-18 | 3.523e-17
34.1 | 6.44e-18 | 3.472e-17
34.2 | 6.14e-18 | 3.442e-17
34.3 | 8.63e-18 | 3.691e-17
34.4 | 8.26e-18 | 3.654e-17
34.5 | 6.68e-18 | 3.496e-17
34.6 | 9.17e-18 | 3.745e-17
34.7 | 7.56e-18 | 3.584e-17
34.8 | 8.94e-18 | 3.722e-17
34.9 | 1.51e-17 | 4.338e-17
35 | 1.03e-17 | 3.858e-17
35.1 | 1.16e-17 | 3.988e-17
35.2 | 8.53e-18 | 3.681e-17
35.3 | 1.29e-17 | 4.118e-17
35.4 | 1.19e-17 | 4.018e-17
35.5 | 8.2e-18 | 3.648e-17
35.6 | 1.02e-17 | 3.848e-17
35.7 | 9.21e-18 | 3.749e-17
35.8 | 9.68e-18 | 3.796e-17
35.9 | 1.44e-17 | 4.268e-17
36 | 1.48e-17 | 4.308e-17
36.1 | 1.77e-17 | 4.598e-17
36.2 | 1.55e-17 | 4.378e-17
36.3 | 1.56e-17 | 4.388e-17
36.4 | 1.32e-17 | 4.148e-17
36.5 | 8.4e-18 | 3.668e-17
36.6 | 4.28e-18 | 3.256e-17
36.7 | 2.93e-18 | 3.121e-17
36.8 | 5.06e-19 | 2.8786e-17
36.9 | 3.78e-18 | 3.206e-17
37 | 5.4e-18 | 3.368e-17
37.1 | 9.88e-18 | 3.816e-17
37.2 | 1.15e-17 | 3.978e-17
37.3 | 1.03e-17 | 3.858e-17
37.4 | 1.09e-17 | 3.918e-17
37.5 | 8.94e-18 | 3.722e-17
37.6 | 6.24e-18 | 3.452e-17
37.7 | 4.59e-18 | 3.287e-17
37.8 | 4.32e-18 | 3.26e-17
37.9 | 8.74e-18 | 3.702e-17
38 | 1.05e-17 | 3.878e-17
38.1 | 1.45e-17 | 4.278e-17
38.2 | 1.82e-17 | 4.648e-17
38.3 | 1.7e-17 | 4.528e-17
38.4 | 1.16e-17 | 3.988e-17
38.5 | 9.21e-18 | 3.749e-17
38.6 | 5.09e-18 | 3.337e-17
38.7 | 3.91e-18 | 3.219e-17
38.8 | 3.07e-18 | 3.135e-17
38.9 | 9.51e-18 | 3.779e-17
39 | 1.28e-17 | 4.108e-17
39.1 | 9.78e-18 | 3.806e-17
39.2 | 1.18e-17 | 4.008e-17
39.3 | 1.66e-17 | 4.488e-17
39.4 | 1.55e-17 | 4.378e-17
39.5 | 1.21e-17 | 4.038e-17
39.6 | 1.17e-17 | 3.998e-17
39.7 | 1.17e-17 | 3.998e-17
39.8 | 7.85e-17 | 1.0678e-16
39.9 | 8.6e-17 | 1.1428e-16
40 | 8.44e-17 | 1.1268e-16
40.1 | 8.59e-17 | 1.1418e-16
40.2 | 8.34e-17 | 1.1168e-16
40.3 | 6.81e-17 | 9.638e-17
40.4 | 4.78e-17 | 7.608e-17
40.5 | -2.8e-17 | 2.8e-19
40.6 | -1.84e-17 | 9.88e-18
40.7 | -6.17e-18 | 2.211e-17
40.8 | 1.28e-18 | 2.956e-17
40.9 | 5.43e-18 | 3.371e-17
41 | 5.83e-18 | 3.411e-17
41.1 | 1.07e-17 | 3.898e-17
41.2 | 7.96e-18 | 3.624e-17
41.3 | 8.03e-18 | 3.631e-17
41.4 | 1.17e-17 | 3.998e-17
41.5 | 8.2e-18 | 3.648e-17
41.6 | 1.06e-17 | 3.888e-17
41.7 | 1e-17 | 3.828e-17
41.8 | 9.11e-18 | 3.739e-17
41.9 | 6.48e-18 | 3.476e-17
42 | 4.28e-18 | 3.256e-17
42.1 | 7.93e-18 | 3.621e-17
42.2 | 1.1e-17 | 3.928e-17
42.3 | 1.15e-17 | 3.978e-17
42.4 | 1.15e-17 | 3.978e-17
42.5 | 9.54e-18 | 3.782e-17
42.6 | 5.94e-18 | 3.422e-17
42.7 | 6.64e-18 | 3.492e-17
42.8 | 7.32e-18 | 3.56e-17
42.9 | 6.41e-18 | 3.469e-17
43 | 5.63e-18 | 3.391e-17
43.1 | 8.13e-18 | 3.641e-17
43.2 | 9.61e-18 | 3.789e-17
43.3 | 1.47e-17 | 4.298e-17
43.4 | 1.4e-17 | 4.228e-17
43.5 | 1.54e-17 | 4.368e-17
43.6 | 7.35e-17 | 1.0178e-16
43.7 | 7.51e-17 | 1.0338e-16
43.8 | 1.33e-16 | 1.6128e-16
43.9 | 2.28e-16 | 2.5628e-16
44 | 3.01e-16 | 3.2928e-16
44.1 | 3.33e-16 | 3.6128e-16
44.2 | 3.32e-16 | 3.6028e-16
44.3 | 3.01e-16 | 3.2928e-16
44.4 | 2.23e-16 | 2.5128e-16
44.5 | 1.18e-16 | 1.4628e-16
44.6 | 6.01e-17 | 8.838e-17
44.7 | -8.57e-18 | 1.971e-17
44.8 | -1.58e-17 | 1.248e-17
44.9 | -5.13e-18 | 2.315e-17
45 | 7.42e-19 | 2.9022e-17
45.1 | 9.78e-19 | 2.9258e-17
45.2 | 2.56e-18 | 3.084e-17
45.3 | 3e-18 | 3.128e-17
45.4 | 3.27e-18 | 3.155e-17
45.5 | 5.77e-18 | 3.405e-17
45.6 | 6.24e-18 | 3.452e-17
45.7 | 4.92e-18 | 3.32e-17
45.8 | 3.41e-18 | 3.169e-17
45.9 | 2.23e-18 | 3.051e-17
46 | 3.98e-18 | 3.226e-17
46.1 | 8.4e-18 | 3.668e-17
46.2 | 7.08e-18 | 3.536e-17
46.3 | 8.23e-18 | 3.651e-17
46.4 | 4.76e-18 | 3.304e-17
46.5 | 3.41e-18 | 3.169e-17
46.6 | 6.54e-18 | 3.482e-17
46.7 | 6.91e-18 | 3.519e-17
46.8 | 5.46e-18 | 3.374e-17
46.9 | 3.54e-18 | 3.182e-17
47 | 4.89e-18 | 3.317e-17
47.1 | 4.49e-18 | 3.277e-17
47.2 | 1.99e-18 | 3.027e-17
47.3 | 3.54e-18 | 3.182e-17
47.4 | 4.45e-18 | 3.273e-17
47.5 | 6.41e-18 | 3.469e-17
47.6 | 1.79e-18 | 3.007e-17
47.7 | 2.5e-18 | 3.078e-17
47.8 | 3.68e-18 | 3.196e-17
47.9 | 4.59e-18 | 3.287e-17
48 | 5.46e-18 | 3.374e-17
48.1 | 1.62e-18 | 2.99e-17
48.2 | 3.07e-18 | 3.135e-17
48.3 | 2.9e-18 | 3.118e-17
48.4 | 3.88e-18 | 3.216e-17
48.5 | -4.38e-19 | 2.7842e-17
48.6 | -2.7e-18 | 2.558e-17
48.7 | 6.07e-19 | 2.8887e-17
48.8 | 2.29e-18 | 3.057e-17
48.9 | 3.78e-18 | 3.206e-17
49 | 3.88e-18 | 3.216e-17
49.1 | 4.25e-18 | 3.253e-17
49.2 | 2.77e-18 | 3.105e-17
49.3 | 2.97e-18 | 3.125e-17
49.4 | 3.71e-18 | 3.199e-17
49.5 | 4.62e-18 | 3.29e-17
49.6 | 5.06e-18 | 3.334e-17
49.7 | 4.92e-18 | 3.32e-17
49.8 | 4.99e-18 | 3.327e-17
49.9 | 2.77e-18 | 3.105e-17
50 | 1.45e-18 | 2.973e-17
50.1 | 1.65e-18 | 2.993e-17
50.2 | 5.87e-18 | 3.415e-17
50.3 | 6.78e-18 | 3.506e-17
50.4 | 7.82e-18 | 3.61e-17
50.5 | 3.78e-18 | 3.206e-17
50.6 | 2.33e-18 | 3.061e-17
50.7 | 3.1e-18 | 3.138e-17
50.8 | -2.36e-19 | 2.8044e-17
50.9 | 6.75e-20 | 2.8348e-17
51 | 4.38e-18 | 3.266e-17
51.1 | 6.54e-18 | 3.482e-17
51.2 | 3.81e-18 | 3.209e-17
51.3 | 4.28e-18 | 3.256e-17
51.4 | 3.98e-18 | 3.226e-17
51.5 | 3.54e-18 | 3.182e-17
51.6 | 6.14e-18 | 3.442e-17
51.7 | 4.99e-18 | 3.327e-17
51.8 | 2.8e-18 | 3.108e-17
51.9 | 1.75e-18 | 3.003e-17
52 | 6.75e-19 | 2.8955e-17
52.1 | 2.87e-18 | 3.115e-17
52.2 | 9.44e-19 | 2.9224e-17
52.3 | -2.7e-18 | 2.558e-17
52.4 | -1.25e-18 | 2.703e-17
52.5 | -2.23e-18 | 2.605e-17
52.6 | -1.96e-18 | 2.632e-17
52.7 | 1.69e-19 | 2.8449e-17
52.8 | 3.04e-19 | 2.8584e-17
52.9 | -1.01e-19 | 2.8179e-17
53 | 3.78e-18 | 3.206e-17
53.1 | 2.9e-18 | 3.118e-17
53.2 | 4.59e-18 | 3.287e-17
53.3 | 6.1e-18 | 3.438e-17
53.4 | 9.88e-18 | 3.816e-17
53.5 | 7.52e-18 | 3.58e-17
53.6 | 4.01e-18 | 3.229e-17
53.7 | 3.84e-18 | 3.212e-17
53.8 | 2.46e-18 | 3.074e-17
53.9 | 5.3e-18 | 3.358e-17
54 | 7.12e-18 | 3.54e-17
54.1 | 4.96e-18 | 3.324e-17
54.2 | 9.44e-19 | 2.9224e-17
54.3 | -1.59e-18 | 2.669e-17
54.4 | 3.71e-19 | 2.8651e-17
54.5 | -5.06e-19 | 2.7774e-17
54.6 | 3.71e-19 | 2.8651e-17
54.7 | 1.08e-18 | 2.936e-17
54.8 | -4.72e-19 | 2.7808e-17
54.9 | 1.69e-19 | 2.8449e-17
55 | 3.71e-19 | 2.8651e-17
55.1 | 2.36e-18 | 3.064e-17
55.2 | 4.89e-18 | 3.317e-17
55.3 | 1.25e-18 | 2.953e-17
55.4 | -2.09e-18 | 2.619e-17
55.5 | 1.69e-19 | 2.8449e-17
55.6 | -2.02e-19 | 2.8078e-17
55.7 | 7.08e-19 | 2.8988e-17
55.8 | 1.21e-18 | 2.949e-17
55.9 | 3.58e-18 | 3.186e-17
56 | 9.44e-19 | 2.9224e-17
56.1 | -9.78e-19 | 2.7302e-17
56.2 | 3.84e-18 | 3.212e-17
56.3 | 6.24e-18 | 3.452e-17
56.4 | 5.97e-18 | 3.425e-17
56.5 | 3.51e-18 | 3.179e-17
56.6 | 2.19e-18 | 3.047e-17
56.7 | 1.59e-18 | 2.987e-17
56.8 | -2.19e-18 | 2.609e-17
56.9 | -1.42e-18 | 2.686e-17
57 | 6.07e-19 | 2.8887e-17
57.1 | -9.44e-19 | 2.7336e-17
57.2 | 1.79e-18 | 3.007e-17
57.3 | 3.61e-18 | 3.189e-17
57.4 | 4.42e-18 | 3.27e-17
57.5 | 4.82e-18 | 3.31e-17
57.6 | 2.9e-18 | 3.118e-17
57.7 | 4.15e-18 | 3.243e-17
57.8 | 6.41e-18 | 3.469e-17
57.9 | 1.75e-18 | 3.003e-17
58 | 7.08e-19 | 2.8988e-17
58.1 | 2.56e-18 | 3.084e-17
58.2 | -4.72e-19 | 2.7808e-17
58.3 | 3.34e-18 | 3.162e-17
58.4 | 2.06e-18 | 3.034e-17
58.5 | 5.77e-18 | 3.405e-17
58.6 | 2.8e-18 | 3.108e-17
58.7 | 7.32e-18 | 3.56e-17
58.8 | 1.01e-17 | 3.838e-17
58.9 | 1.09e-17 | 3.918e-17
59 | 3.47e-18 | 3.175e-17
59.1 | 1.05e-18 | 2.933e-17
59.2 | 1.18e-18 | 2.946e-17
59.3 | 1.48e-18 | 2.976e-17
59.4 | 4.52e-18 | 3.28e-17
59.5 | 4.08e-18 | 3.236e-17
59.6 | 3.78e-18 | 3.206e-17
59.7 | 5.33e-18 | 3.361e-17
59.8 | 6.78e-18 | 3.506e-17
59.9 | 8.6e-18 | 3.688e-17
60 | 2.66e-18 | 3.094e-17
60.1 | -7.76e-19 | 2.7504e-17
60.2 | -1.42e-18 | 2.686e-17
60.3 | 5.73e-19 | 2.8853e-17
60.4 | 2.6e-18 | 3.088e-17
60.5 | 4.28e-18 | 3.256e-17
60.6 | 4.99e-18 | 3.327e-17
60.7 | 2.77e-18 | 3.105e-17
60.8 | -9.78e-19 | 2.7302e-17
60.9 | 1.55e-18 | 2.983e-17
61 | 4.11e-18 | 3.239e-17
61.1 | 3.51e-18 | 3.179e-17
61.2 | 1.75e-18 | 3.003e-17
61.3 | 2.43e-18 | 3.071e-17
61.4 | 1.79e-18 | 3.007e-17
61.5 | -1.35e-19 | 2.8145e-17
61.6 | 1.08e-18 | 2.936e-17
61.7 | 3.81e-18 | 3.209e-17
61.8 | 4.72e-18 | 3.3e-17
61.9 | 5.43e-18 | 3.371e-17
62 | 6.17e-18 | 3.445e-17
62.1 | 5.43e-18 | 3.371e-17
62.2 | 1.32e-18 | 2.96e-17
62.3 | 7.08e-19 | 2.8988e-17
62.4 | 2.77e-18 | 3.105e-17
62.5 | 8.09e-19 | 2.9089e-17
62.6 | 5.4e-19 | 2.882e-17
62.7 | -4.05e-19 | 2.7875e-17
62.8 | 5.06e-19 | 2.8786e-17
62.9 | 7.76e-19 | 2.9056e-17
63 | 1.69e-19 | 2.8449e-17
63.1 | 2.63e-18 | 3.091e-17
63.2 | 2.39e-18 | 3.067e-17
63.3 | 3.54e-18 | 3.182e-17
63.4 | 4.52e-18 | 3.28e-17
63.5 | 3.41e-18 | 3.169e-17
63.6 | 4.76e-18 | 3.304e-17
63.7 | 4.76e-18 | 3.304e-17
63.8 | 5.16e-18 | 3.344e-17
63.9 | 6e-18 | 3.428e-17
64 | 3.68e-18 | 3.196e-17
64.1 | 5.7e-18 | 3.398e-17
64.2 | 4.49e-18 | 3.277e-17
64.3 | 6.78e-18 | 3.506e-17
64.4 | 3.47e-18 | 3.175e-17
64.5 | 3.64e-18 | 3.192e-17
64.6 | 4.59e-18 | 3.287e-17
64.7 | 1.69e-18 | 2.997e-17
64.8 | 4.89e-18 | 3.317e-17
64.9 | 6.48e-18 | 3.476e-17
65 | 4.99e-18 | 3.327e-17
65.1 | 3.58e-18 | 3.186e-17
65.2 | 3.84e-18 | 3.212e-17
65.3 | 2.83e-18 | 3.111e-17
65.4 | 4.32e-18 | 3.26e-17
65.5 | 5.13e-18 | 3.341e-17
65.6 | 3.71e-18 | 3.199e-17
65.7 | 4.08e-18 | 3.236e-17
65.8 | 3.58e-18 | 3.186e-17
65.9 | 6.85e-18 | 3.513e-17
66 | 8.84e-18 | 3.712e-17
66.1 | 9.34e-18 | 3.762e-17
66.2 | 6.54e-18 | 3.482e-17
66.3 | 1.05e-17 | 3.878e-17
66.4 | 7.89e-18 | 3.617e-17
66.5 | 5.7e-18 | 3.398e-17
66.6 | 6.17e-18 | 3.445e-17
66.7 | 4.18e-18 | 3.246e-17
66.8 | 2.23e-18 | 3.051e-17
66.9 | 1.75e-18 | 3.003e-17
67 | 5.9e-18 | 3.418e-17
67.1 | 3.27e-18 | 3.155e-17
67.2 | 7.76e-19 | 2.9056e-17
67.3 | 1.18e-18 | 2.946e-17
67.4 | -1.69e-19 | 2.8111e-17
67.5 | 3.17e-18 | 3.145e-17
67.6 | 0 | 2.828e-17
67.7 | 1.11e-18 | 2.939e-17
67.8 | 3.04e-18 | 3.132e-17
67.9 | 1.52e-18 | 2.98e-17
68 | 2.33e-18 | 3.061e-17
68.1 | 1.79e-18 | 3.007e-17
68.2 | 1.55e-18 | 2.983e-17
68.3 | -5.73e-19 | 2.7707e-17
68.4 | 8.43e-19 | 2.9123e-17
68.5 | 1.69e-19 | 2.8449e-17
68.6 | -2.23e-18 | 2.605e-17
68.7 | -3.64e-18 | 2.464e-17
68.8 | -1.21e-18 | 2.707e-17
68.9 | -1.38e-18 | 2.69e-17
69 | -3.41e-18 | 2.487e-17
69.1 | 8.77e-19 | 2.9157e-17
69.2 | 1.48e-18 | 2.976e-17
69.3 | -6.41e-19 | 2.7639e-17
69.4 | -1.72e-18 | 2.656e-17
69.5 | 3.37e-20 | 2.8314e-17
69.6 | 2.56e-18 | 3.084e-17
69.7 | 1.59e-18 | 2.987e-17
69.8 | 2.19e-18 | 3.047e-17
69.9 | 1.92e-18 | 3.02e-17
70 | 4.72e-19 | 2.8752e-17
70.1 | 1.01e-19 | 2.8381e-17
70.2 | 6.41e-19 | 2.8921e-17
70.3 | 5.16e-18 | 3.344e-17
70.4 | 4.65e-18 | 3.293e-17
70.5 | 1.82e-18 | 3.01e-17
70.6 | 2.5e-18 | 3.078e-17
70.7 | 2.5e-18 | 3.078e-17
70.8 | 2.77e-18 | 3.105e-17
70.9 | 4.08e-18 | 3.236e-17
71 | 4.49e-18 | 3.277e-17
71.1 | 6.04e-18 | 3.432e-17
71.2 | 4.92e-18 | 3.32e-17
71.3 | 1.92e-18 | 3.02e-17
71.4 | 3.37e-20 | 2.8314e-17
71.5 | 9.78e-19 | 2.9258e-17
71.6 | 6e-18 | 3.428e-17
71.7 | 5.36e-18 | 3.364e-17
71.8 | 8.43e-19 | 2.9123e-17
71.9 | 3.95e-18 | 3.223e-17
72 | 3.81e-18 | 3.209e-17
72.1 | 4.05e-18 | 3.233e-17
72.2 | 2.39e-18 | 3.067e-17
72.3 | 4.32e-18 | 3.26e-17
72.4 | 3.91e-18 | 3.219e-17
72.5 | 3.71e-18 | 3.199e-17
72.6 | 4.59e-18 | 3.287e-17
72.7 | 7.25e-18 | 3.553e-17
72.8 | 5.57e-18 | 3.385e-17
72.9 | 8.94e-18 | 3.722e-17
73 | 7.22e-18 | 3.55e-17
73.1 | 2.7e-18 | 3.098e-17
73.2 | 8.77e-19 | 2.9157e-17
73.3 | 1.72e-18 | 3e-17
73.4 | 2.66e-18 | 3.094e-17
73.5 | 3.64e-18 | 3.192e-17
73.6 | 6.61e-18 | 3.489e-17
73.7 | 1.38e-18 | 2.966e-17
73.8 | 2.53e-18 | 3.081e-17
73.9 | 2.16e-18 | 3.044e-17
74 | 3.61e-18 | 3.189e-17
74.1 | 1.62e-18 | 2.99e-17
74.2 | 1.38e-18 | 2.966e-17
74.3 | 2.7e-19 | 2.855e-17
74.4 | 2.23e-18 | 3.051e-17
74.5 | 1.01e-18 | 2.929e-17
74.6 | -1.01e-19 | 2.8179e-17
74.7 | 1.01e-19 | 2.8381e-17
74.8 | 1.69e-18 | 2.997e-17
74.9 | 3.37e-20 | 2.8314e-17
75 | 2.33e-18 | 3.061e-17
75.1 | 1.92e-18 | 3.02e-17
75.2 | 3.2e-18 | 3.148e-17
75.3 | 2.43e-18 | 3.071e-17
75.4 | 4.01e-18 | 3.229e-17
75.5 | 2.56e-18 | 3.084e-17
75.6 | 1.45e-18 | 2.973e-17
75.7 | 1.35e-18 | 2.963e-17
75.8 | 1.35e-18 | 2.963e-17
75.9 | 9.11e-19 | 2.9191e-17
76 | 2.36e-18 | 3.064e-17
76.1 | -7.42e-19 | 2.7538e-17
76.2 | 8.77e-19 | 2.9157e-17
76.3 | 1.69e-18 | 2.997e-17
76.4 | -4.38e-19 | 2.7842e-17
76.5 | -8.77e-19 | 2.7403e-17
76.6 | 2.26e-18 | 3.054e-17
76.7 | 3.24e-18 | 3.152e-17
76.8 | 1.69e-18 | 2.997e-17
76.9 | 2.02e-19 | 2.8482e-17
77 | -9.44e-19 | 2.7336e-17
77.1 | 2.66e-18 | 3.094e-17
77.2 | 3.47e-18 | 3.175e-17
77.3 | 4.72e-19 | 2.8752e-17
77.4 | 2.97e-18 | 3.125e-17
77.5 | 3.47e-18 | 3.175e-17
77.6 | 5.03e-18 | 3.331e-17
77.7 | 1.21e-18 | 2.949e-17
77.8 | 0 | 2.828e-17
77.9 | 3.37e-19 | 2.8617e-17
78 | 1.08e-18 | 2.936e-17
78.1 | 1.59e-18 | 2.987e-17
78.2 | 4.05e-18 | 3.233e-17
78.3 | 1.79e-18 | 3.007e-17
78.4 | 4.55e-18 | 3.283e-17
78.5 | 4.38e-18 | 3.266e-17
78.6 | 2.9e-18 | 3.118e-17
78.7 | -5.73e-19 | 2.7707e-17
78.8 | 3.24e-18 | 3.152e-17
78.9 | 2.73e-18 | 3.101e-17
79 | 4.65e-18 | 3.293e-17
79.1 | 6.37e-18 | 3.465e-17
79.2 | 6.95e-18 | 3.523e-17
79.3 | 3.14e-18 | 3.142e-17
79.4 | 5.6e-18 | 3.388e-17
79.5 | -3.37e-19 | 2.7943e-17
79.6 | 7.76e-19 | 2.9056e-17
79.7 | 1.92e-18 | 3.02e-17
79.8 | 4.32e-18 | 3.26e-17
79.9 | 3.24e-18 | 3.152e-17
80 | 4.69e-18 | 3.297e-17
80.1 | 1.35e-19 | 2.8415e-17
80.2 | 1.69e-19 | 2.8449e-17
80.3 | 4.01e-18 | 3.229e-17
80.4 | 1.79e-18 | 3.007e-17
80.5 | -3.37e-20 | 2.8246e-17
80.6 | -1.11e-18 | 2.717e-17
80.7 | 1.55e-18 | 2.983e-17
80.8 | 8.77e-19 | 2.9157e-17
80.9 | -1.82e-18 | 2.646e-17
81 | 2.19e-18 | 3.047e-17
81.1 | 2.16e-18 | 3.044e-17
81.2 | 2.16e-18 | 3.044e-17
81.3 | 3.07e-18 | 3.135e-17
81.4 | 5.6e-18 | 3.388e-17
81.5 | 7.29e-18 | 3.557e-17
81.6 | 7.76e-18 | 3.604e-17
81.7 | 4.79e-18 | 3.307e-17
81.8 | 3.24e-18 | 3.152e-17
81.9 | -6.75e-20 | 2.8212e-17
82 | 2.6e-18 | 3.088e-17
82.1 | 3.47e-18 | 3.175e-17
82.2 | 1.65e-18 | 2.993e-17
82.3 | -7.76e-19 | 2.7504e-17
82.4 | -3.37e-19 | 2.7943e-17
82.5 | 8.43e-19 | 2.9123e-17
82.6 | -3.71e-19 | 2.7909e-17
82.7 | 2.16e-18 | 3.044e-17
82.8 | 2.87e-18 | 3.115e-17
82.9 | 1.65e-18 | 2.993e-17
83 | -1.32e-18 | 2.696e-17
83.1 | 8.77e-19 | 2.9157e-17
83.2 | 2.63e-18 | 3.091e-17
83.3 | 3.07e-18 | 3.135e-17
83.4 | 3.07e-18 | 3.135e-17
83.5 | -2.02e-19 | 2.8078e-17
83.6 | -3.61e-18 | 2.467e-17
83.7 | -1.15e-18 | 2.713e-17
83.8 | 1.21e-18 | 2.949e-17
83.9 | 2.8e-18 | 3.108e-17
84 | -1.86e-18 | 2.642e-17
84.1 | -7.76e-19 | 2.7504e-17
84.2 | 2.09e-18 | 3.037e-17
84.3 | 4.25e-18 | 3.253e-17
84.4 | 4.28e-18 | 3.256e-17
84.5 | 1.96e-18 | 3.024e-17
84.6 | 4.69e-18 | 3.297e-17
84.7 | 1.99e-18 | 3.027e-17
84.8 | 3.07e-18 | 3.135e-17
84.9 | -1.21e-18 | 2.707e-17
85 | -6.07e-19 | 2.7673e-17
85.1 | -3.37e-19 | 2.7943e-17
85.2 | -1.18e-18 | 2.71e-17
85.3 | -1.75e-18 | 2.653e-17
85.4 | -3.64e-18 | 2.464e-17
85.5 | 2.02e-18 | 3.03e-17
85.6 | 5.73e-19 | 2.8853e-17
85.7 | -4.72e-19 | 2.7808e-17
85.8 | 4.05e-18 | 3.233e-17
85.9 | 2.8e-18 | 3.108e-17
86 | 3.95e-18 | 3.223e-17
86.1 | 6.81e-18 | 3.509e-17
86.2 | 4.72e-18 | 3.3e-17
86.3 | 5.9e-18 | 3.418e-17
86.4 | 3.41e-18 | 3.169e-17
86.5 | 3.27e-18 | 3.155e-17
86.6 | 4.05e-18 | 3.233e-17
86.7 | 4.11e-18 | 3.239e-17
86.8 | 2.26e-18 | 3.054e-17
86.9 | 9.44e-19 | 2.9224e-17
87 | 9.78e-19 | 2.9258e-17
87.1 | 4.32e-18 | 3.26e-17
87.2 | 3.31e-18 | 3.159e-17
87.3 | 4.11e-18 | 3.239e-17
87.4 | 6.41e-19 | 2.8921e-17
87.5 | -2.66e-18 | 2.562e-17
87.6 | -3.31e-18 | 2.497e-17
87.7 | -4.72e-19 | 2.7808e-17
87.8 | -1.89e-18 | 2.639e-17
87.9 | 3.37e-19 | 2.8617e-17
88 | 1.72e-18 | 3e-17
88.1 | -6.07e-19 | 2.7673e-17
88.2 | 0 | 2.828e-17
88.3 | 4.01e-18 | 3.229e-17
88.4 | 5.16e-18 | 3.344e-17
88.5 | 1.96e-18 | 3.024e-17
88.6 | 4.15e-18 | 3.243e-17
88.7 | 2.6e-18 | 3.088e-17
88.8 | 1.82e-18 | 3.01e-17
88.9 | 5.03e-18 | 3.331e-17
89 | 4.92e-18 | 3.32e-17
89.1 | 6.75e-20 | 2.8348e-17
89.2 | -8.43e-19 | 2.7437e-17
89.3 | -4.55e-18 | 2.373e-17
89.4 | -3.37e-18 | 2.491e-17
89.5 | -1.75e-18 | 2.653e-17
89.6 | -6.75e-19 | 2.7605e-17
89.7 | 2.36e-18 | 3.064e-17
89.8 | 4.92e-18 | 3.32e-17
89.9 | 6.48e-18 | 3.476e-17
90 | 2.02e-18 | 3.03e-17
90.1 | 1.69e-19 | 2.8449e-17
90.2 | 2.36e-18 | 3.064e-17
90.3 | 4.55e-18 | 3.283e-17
90.4 | 4.18e-18 | 3.246e-17
90.5 | 2.77e-18 | 3.105e-17
90.6 | 2.97e-18 | 3.125e-17
90.7 | 3.1e-18 | 3.138e-17
90.8 | 4.89e-18 | 3.317e-17
90.9 | 5.26e-18 | 3.354e-17
91 | 2.8e-18 | 3.108e-17
91.1 | 2.8e-18 | 3.108e-17
91.2 | 2.66e-18 | 3.094e-17
91.3 | 3.34e-18 | 3.162e-17
91.4 | 1.65e-18 | 2.993e-17
91.5 | 1.25e-18 | 2.953e-17
91.6 | 3.07e-18 | 3.135e-17
91.7 | 3.07e-18 | 3.135e-17
91.8 | 2.63e-18 | 3.091e-17
91.9 | -6.41e-19 | 2.7639e-17
92 | 1.99e-18 | 3.027e-17
92.1 | 6.07e-18 | 3.435e-17
92.2 | 3.58e-18 | 3.186e-17
92.3 | 4.45e-18 | 3.273e-17
92.4 | 3.91e-18 | 3.219e-17
92.5 | 2.8e-18 | 3.108e-17
92.6 | 3.64e-18 | 3.192e-17
92.7 | 9.78e-19 | 2.9258e-17
92.8 | 1.15e-18 | 2.943e-17
92.9 | 1.62e-18 | 2.99e-17
93 | -4.38e-19 | 2.7842e-17
93.1 | -4.38e-19 | 2.7842e-17
93.2 | 4.32e-18 | 3.26e-17
93.3 | 2.97e-18 | 3.125e-17
93.4 | 1.52e-18 | 2.98e-17
93.5 | -1.92e-18 | 2.636e-17
93.6 | 7.76e-19 | 2.9056e-17
93.7 | 1.48e-18 | 2.976e-17
93.8 | -3.04e-19 | 2.7976e-17
93.9 | 9.44e-19 | 2.9224e-17
94 | 3.37e-20 | 2.8314e-17
94.1 | 6.41e-19 | 2.8921e-17
94.2 | 1.55e-18 | 2.983e-17
94.3 | 4.11e-18 | 3.239e-17
94.4 | 3.27e-18 | 3.155e-17
94.5 | 0 | 2.828e-17
94.6 | 3.04e-19 | 2.8584e-17
94.7 | 3.1e-18 | 3.138e-17
94.8 | 4.76e-18 | 3.304e-17
94.9 | 3.14e-18 | 3.142e-17
95 | 7.42e-19 | 2.9022e-17
95.1 | 2.56e-18 | 3.084e-17
95.2 | 7.42e-19 | 2.9022e-17
95.3 | 2.7e-19 | 2.855e-17
95.4 | 5.73e-19 | 2.8853e-17
95.5 | 1.48e-18 | 2.976e-17
95.6 | 2.66e-18 | 3.094e-17
95.7 | 7.08e-19 | 2.8988e-17
95.8 | 1.62e-18 | 2.99e-17
95.9 | 4.05e-18 | 3.233e-17
96 | 3.98e-18 | 3.226e-17
96.1 | 2.8e-18 | 3.108e-17
96.2 | 2.56e-18 | 3.084e-17
96.3 | 1.01e-18 | 2.929e-17
96.4 | 2.7e-18 | 3.098e-17
96.5 | 5.73e-18 | 3.401e-17
96.6 | 6.68e-18 | 3.496e-17
96.7 | 4.55e-18 | 3.283e-17
96.8 | 5.97e-18 | 3.425e-17
96.9 | 4.89e-18 | 3.317e-17
97 | 3e-18 | 3.128e-17
97.1 | 2.83e-18 | 3.111e-17
97.2 | 5.06e-19 | 2.8786e-17
97.3 | 7.08e-19 | 2.8988e-17
97.4 | -1.28e-18 | 2.7e-17
97.5 | 2.7e-19 | 2.855e-17
97.6 | 9.78e-19 | 2.9258e-17
97.7 | -1.01e-18 | 2.727e-17
97.8 | -3.88e-18 | 2.44e-17
97.9 | -3.78e-18 | 2.45e-17
98 | 1.89e-18 | 3.017e-17
98.1 | -1.35e-19 | 2.8145e-17
98.2 | -2.19e-18 | 2.609e-17
98.3 | 1.99e-18 | 3.027e-17
98.4 | 3.14e-18 | 3.142e-17
98.5 | 6.24e-18 | 3.452e-17
98.6 | 5.3e-18 | 3.358e-17
98.7 | 4.32e-18 | 3.26e-17
98.8 | 1.35e-18 | 2.963e-17
98.9 | 1.75e-18 | 3.003e-17
99 | 1.08e-18 | 2.936e-17
99.1 | 5.13e-18 | 3.341e-17
99.2 | 4.62e-18 | 3.29e-17
99.3 | 8.77e-19 | 2.9157e-17
99.4 | -1.08e-18 | 2.72e-17
99.5 | -1.48e-18 | 2.68e-17
99.6 | -6.07e-19 | 2.7673e-17
99.7 | 8.77e-19 | 2.9157e-17
99.8 | 9.44e-19 | 2.9224e-17
99.9 | 3.58e-18 | 3.186e-17
100 | 2.39e-18 | 3.067e-17

## F82310-AOSEMs-393-cal
2010-08-23 | 2010 04:35:50 PM
Residual Gas Analyzer Software
RGA Software Version | 3.205.003
Pressure vs Time Scan Setup:
Active channels in Scan | 1
Units | Amps
Sample Period | 1 | sec
Focus Voltage | 90 | Volts
Ion Energy | HIGH
Electron Energy | 70 | eV
CEM Voltage | 1000 | Volts
CEM Gain | 1000000
Sensitivity Factor | 0.0001
Filament Current | 1 | mAmps
Focus Voltage | 90 | Volts
Ion Energy | HIGH
Electron Energy | 70 | eV
CEM Voltage | 1000 | Volts
CEM Gain | 1000000
Sensitivity Factor | 0.0001
Filament Current | 1 | mAmps
Start time | Aug 23 | 2010 04:35:50 PM
Channel | Mass(amu) | Name | Cal Factor | Noise Floor | CEM Status
1 40.00 Argon 0.70 3 OFF
Time(s) Channel#1
0 | 9.8523e-17
1 | 2.1532e-16
2 | 5.0383e-17
3 | 5.2696e-17
4 | 5.9709e-17
5 | 1.8581e-16
6 | 5.3074e-17
7 | 5.6946e-17
8 | 2.0531e-16
9 | 5.0029e-17
10 | 2.1777e-16
11 | 2.4537e-16
12 | 2.1173e-16
13 | 4.8943e-17
14 | 4.6983e-17
15 | 5.4845e-17
16 | 2.2002e-16
17 | 1.0853e-12
18 | 2.4106e-13
19 | 3.0945e-11
20 | 3.0945e-11
21 | 3.0945e-11
22 | 1.8445e-11
23 | 2.9585e-12
24 | 6.7371e-13
25 | 1.6459e-13
26 | 6.1988e-14
27 | 4.0893e-14
28 | 3.5191e-14
29 | 3.2425e-14
30 | 3.0956e-14
31 | 2.9976e-14
32 | 2.8979e-14
33 | 2.7833e-14
34 | 2.7164e-14
35 | 2.6707e-14
36 | 2.6204e-14
37 | 2.5437e-14
38 | 2.5312e-14
39 | 2.4883e-14
40 | 2.4531e-14
41 | 2.4285e-14
42 | 2.3734e-14
43 | 2.3429e-14
44 | 2.3061e-14
45 | 2.2805e-14
46 | 2.2894e-14
47 | 2.268e-14
48 | 2.2163e-14
49 | 2.2239e-14
50 | 2.2188e-14
51 | 2.1801e-14
52 | 2.2036e-14
53 | 2.1534e-14
54 | 2.1447e-14
55 | 2.1112e-14
56 | 2.1032e-14
57 | 2.1126e-14
58 | 2.1152e-14
59 | 2.0876e-14
60 | 2.0784e-14
61 | 2.0937e-14
62 | 2.0505e-14
63 | 2.0397e-14
64 | 2.0482e-14
65 | 2.0554e-14
66 | 2.0296e-14
67 | 2.0096e-14
68 | 2.0486e-14
69 | 2.0077e-14
70 | 1.9963e-14
71 | 1.9899e-14
72 | 1.9967e-14
73 | 2.0065e-14
74 | 1.8857e-14
75 | 5.4053e-15
76 | 2.5633e-15
77 | 2.0282e-15
78 | 1.8663e-15
79 | 1.841e-15
80 | 1.7967e-15
81 | 1.7941e-15
82 | 1.749e-15
83 | 1.7515e-15
84 | 1.7294e-15
85 | 1.6587e-15
86 | 1.6521e-15
87 | 1.5991e-15
88 | 1.6072e-15
89 | 1.5511e-15
90 | 1.5454e-15
91 | 1.5373e-15
92 | 1.4574e-15
93 | 1.5054e-15
94 | 1.4451e-15
95 | 1.4445e-15
96 | 1.4518e-15
97 | 1.3999e-15
98 | 1.4006e-15
99 | 1.351e-15
100 | 1.3654e-15
101 | 1.383e-15
102 | 1.3137e-15
103 | 1.2991e-15
104 | 1.2467e-15
105 | 1.2838e-15
106 | 1.2519e-15
107 | 1.2607e-15
108 | 1.2389e-15
109 | 1.201e-15
110 | 1.1869e-15
111 | 1.1994e-15
112 | 1.2068e-15
113 | 1.1718e-15
114 | 1.1792e-15
115 | 1.1754e-15
116 | 1.1519e-15
117 | 1.1784e-15
118 | 1.1109e-15
119 | 1.1073e-15
120 | 1.1028e-15
121 | 1.0984e-15
122 | 1.07e-15
123 | 1.0761e-15
124 | 1.0528e-15
125 | 1.0462e-15
126 | 1.04e-15
127 | 1.0141e-15
128 | 1.0183e-15
129 | 1.0222e-15
130 | 1.028e-15
131 | 1.0021e-15
132 | 9.7524e-16
133 | 9.5302e-16
134 | 9.9358e-16
135 | 9.768e-16

## Calibration
Calibration with leak
Calib leak rate | 2.36e-10 | torr l/s | (Argon)
pump speed | ? | l/s | oven F
AMU 40 (w/leak open) | 2e-14 | amps
AMU 40 (background) | 1e-16 | amps
Calib leak contributes | 1.99e-14 | amps | = (w/leak open) - (background)
Leak Rate Calibration factor | 11859.2965 | torr l/s/A
Pressure Calibration factor | #VALUE! | torr/A
Leak Rate Cal factor per AOSEM | 29.3547 | torr l/s/A/AOSEM
The load consists of 404 A-OSEM bodies (D0901048-v4)
| 31879.64 | cm^2 | for the test
The A-OSEM body is D0901048-v4
| 78.91 | cm^2 | surface area of each A-OSEM head
| 104 | units | in the vertex volume
| 112 | units | in the LHO diagonal volume
| 8 | units | in the LHO end station volume