## how to use this workbook
How to use this workbook:
1) Get an E-number for this test report from https://dcc.ligo.org
a) Label the document "Optical Contamination Test Results: <name>", where <name> is a descriptive name of the article under test
b) For the E-<number>-X0 version DCC entry, give a COMPLETE description of the article(s) under test in the Abstract (not the "Changes and Notes"), including manufacturer/source, material grade or assembly/device model number, and quantity/amount. For example, it is not sufficient to say "photodiode", but rather state the manufacturer's name and model number, such as "Vishay TSTS7100 IR-LED". Likewise "PEEK" is not a sufficient material descriptor, but "Victrex PEEK grade 450GL" is a complete description.
c) For the E-<number>-X0 version DCC entry, give a URL link to the location where intermediate test results will be posted in the "Notes and Changes" field.
2) Rename this template workbook to "E<number>-v1 <name>.xls"
3) Fill in only the light yellow colored cells in the sheet labelled "scaled Loss rate".
You must decide the appropriate units to characterize the material, or assembly or device under test. For example, for a device or assembly the appropriate units is likely "each", whereas for a bulk material (say PEEK) the appropriate unit is likely surface area (cm^2). For a lubricant, such as Krytox, the appropriate units may be mass, or volume.
4) Fill in only the light yellow colored cells in the sheets labelled "d1", "d2", "d3", etc.
Each sheet is a data set taken at a particular day and time. In the first data set sheet (d1), the start date of the test and the cavity number is entered. Data for the TEM 00-01 beat frequency as a function of power and the ring down data are entered for each data set. The data is plotted and fitted automatically. The sheet labelled "data" summarizes the absorption and total loss measurements as a function of time in a tabular listing. Sheets "absorption" and "total loss" summarize the absortion and total loss in plots and a least squares ft, as the data is entered. The "scaled loss rate" sheet summarizes the results scaled to the LIGO Observatory for the intended use and is updated automatically as data is entered.
5) The test is terminated when a determination is made that the material/assembly/device under test has either passed or failed. At this point, use the "Convert to Adobe PDF" tool to create a file with just the spredsheets labelled "scaled loss rate", "absorption" and "total loss" which is named "E<number>-v1 <name>.pdf".
6) Select "Create E<number>-v1" in the DCC entry for this document and upload the pdf file "E<number>-v1 <name>.pdf" as the "Main file in document" and this Excel file "E<number>-v1 <name>.xls" as an "Other file" (not Main). Then in either the abstract or the "Notes and Changes" field indicate the test results (passed or failed).
7) File any other documents, which help to define the material/assembly/device which was tested, as "Other files" under the DCC entry for E<number>-v1, such as photos or descriptions of the test (if it was a unique test, such as taken at elevated temperatures).
8) File proprietary data (data not owned by LIGO Lab, even if posted to the web) about the material/assembly/device such as MSDS sheets, specifications, etc. This data should be filed under a separate E-number and then linked from the E-number which corresponds to the optical contamination cavity test results. Even if the MSDS, specifications, etc. are already on the web, they should be filed in the DCC, in case the material/assembly/device is discontinued (in which case the manufacturer often removes this data from their web sites.)

## scaled loss rate
LIGO E-Document Number: |  |  | E1000453-v1
Sample Test:
| Material under test: |  | EP30-2 Bonded Quartz Glass Disks
| units |  |  | 50.6707 | cm^2 bond surface
| absorption | 0.88613 | ± | 0.17088 | ppm/yr | 1 sigma
| scatter | 20.1188 | ± | 2.8525 | ppm/yr | 1 sigma
| max. normalized absorption |  |  | 0.024233 | ppm/yr/unit | 2 sigma
| max. normalized scatter |  |  | 0.50964 | ppm/yr/unit | 2 sigma
| test turbopump speed (liter/s) |  |  | 24.3916 | torr/liter/sec
Scaled to LIGO:
| LIGO Vacuum Volume |  |  | Vertex | LHO Diagonal | End | Comments
| Quantity (units) |  |  | 60 | 77 | 11 | see E1000386-v1 for an estimate of amounts
| LIGO ion pumping speed (liter/s) |  |  | 6800 | 6800 | 1700 | see E0900398 or PSI V049-1-078 for pump rates
| pumping speed ratio (test/LIGO) |  |  | 0.003587 | 0.003587 | 0.014348 | does not include cryo-pump and effective pumping from the Beam Tube
| max. absorption (ppm/yr) |  |  | 0.0052154 | 0.0066931 | 0.0038246 | * Limit is < 0.02 ppm/yr for a single source
| max. scatter (ppm/yr) |  |  | 0.10968 | 0.14076 | 0.080435 | * Limit is < 0.2 ppm/yr for a single source
* The overall limit on contamination loss on optics for AdL is < 0.5 ppm/yr absorption and < 4 ppm/yr scatter from all sources, per Table 4 of the COC Design Requirements Document (T000127-v1). It is assumed that ~20 sources could contribute.
10 pairs of quartz disks each 2.54 cm diameter are bonded together with a thin coating of Masterbond EP30-2 Epoxy.
Epoxy was prepared with 10 parts A to 1 part B per manufactures instruction sheet and cured at room temperature for 24 Hrs. in the clean room.
then baked at 100F for 4 Hrs. for max. hardness.
| 50.6707 | cm^2 bonded surface (not directly exposed to the vacuum)
|  |  |  | 2.75 | in, dia, flex hose
|  |  |  | 36 | in, length, flex hose
|  |  |  | 6.985 | cm, dia
|  |  |  | 91.44 | cm, length
|  |  |  | 463.669 | C1, aperture, liter/sec
|  |  |  | 45.0971 | C2, flex hose, liter/sec
|  |  |  | 41.0997 | Ct, liter/sec
|  |  |  | 60 | S, liter/sec [Varian Dry TPS-compact with TV81M]
|  |  |  | 24.3916 | Seff, liter/sec

## absorption
LIGO E-Document Number: |  | E1000453-v1
Test Material/Assy./Device: |  | EP30-2 Bonded Quartz Glass Disks
Absorption fitting
| Slope | 0.00010116 | 1.2889 | Y-intercept
| Standard Error | 1.9507e-05 | 0.024055 | Standard error
| r2 | 0.45661 | 0.069173 | sey
| F | 26.89 | 32 | df
| ssreg | 0.12867 | 0.15312 | ssresid
| Absorption change rate (ppm/yr) |  | ± sigma (ppm/yr)
| 0.88613 |  | 0.17088

## total loss
LIGO E-Document Number: |  | E1000453-v1
Test Material/Assy./Device: |  | EP30-2 Bonded Quartz Glass Disks
Total loss fitting
| Slope | 0.0022967 | 137.8913 | Y-intercept
| Standard Error | 0.00032563 | 0.40154 | Standard error
| r2 | 0.60854 | 1.1547 | sey
| F | 49.745 | 32 | df
| ssreg | 66.3242 | 42.6651 | ssresid
| Total loss change rate (ppm/yr) |  | ± sigma (ppm/yr)
| 20.1188 |  | 2.8525

## data
| Time (hr) | Absorption (ppm) | Total Loss (ppm) | EP30-2 Bonded Quartz Glass Disks
1 | 57 | 1.3282 | 138.602
2 | 129.5 | 1.2965 | 138.4337
3 | 177 | 1.3445 | 138.1574
4 | 226.5 | 1.2658 | 138.3977
5 | 299 | 1.2768 | 138.8379
6 | 352 | 1.3018 | 138.3809
7 | 384 | 1.5831 | 143.1699
8 | 537.5 | 1.2969 | 137.6623
9 | 560.5 | 1.3134 | 138.3141
10 | 657 | 1.3032 | 139.1013
11 | 682 | 1.3157 | 138.8712
12 | 729 | 1.2947 | 138.1874
13 | 802 | 1.3105 | 138.5863
14 | 848.5 | 1.3502 | 139.1008
15 | 898 | 1.2815 | 138.5319
16 | 992 | 1.3663 | 138.6509
17 | 1016 | 1.3592 | 139.2678
18 | 1064 | 1.438 | 141.5985
19 | 1136 | 1.4696 | 141.3542
20 | 1184.5 | 1.4703 | 141.5934
21 | 1236 | 1.4609 | 141.885
22 | 1356 | 1.4282 | 140.739 | why must this line be hard coded?
23 | 1400.5 | 1.335 | 139.8396
24 | 1474 | 1.5721 | 142.2918
25 | 1520.5 | 1.4628 | 142.4779
26 | 1641 | 1.5343 | 142.482
27 | 1689 | 1.4779 | 142.3602
28 | 1736 | 1.4424 | 142.3027
29 | 1814 | 1.4611 | 141.6318
30 | 1857 | 1.5192 | 142.0781
31 | 1906 | 1.4784 | 141.9473
32 | 2001.5 | 1.4666 | 142.6847
33 | 2031 | 1.4583 | 142.4781
34 | 2078 | 1.4476 | 142.0712
35 | #N/A | #N/A | #N/A
36 | #N/A | #N/A | #N/A
37 | #N/A | #N/A | #N/A
38 | #N/A | #N/A | #N/A
39 | #N/A | #N/A | #N/A
40 | #N/A | #N/A | #N/A
41 | #N/A | #N/A | #N/A
42 | #N/A | #N/A | #N/A
43 | #N/A | #N/A | #N/A
44 | #N/A | #N/A | #N/A
45 | #N/A | #N/A | #N/A
46 | #N/A | #N/A | #N/A
47 | #N/A | #N/A | #N/A
48 | #N/A | #N/A | #N/A
49 | #N/A | #N/A | #N/A
50 | #N/A | #N/A | #N/A

## d1
Cavity No | Sample | Comment
1 | EP30-2 Quartz Disks | Perkin Elmer Silicon Quadrant Diode, 8mm Active Diameter
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-09-24 | 9 | 2 | 57 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
24.5 | 1.2e-08 | 24.9 | 44.7504
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
60 | 0.589 | 0.149 | 74.55 | 2.405e-05 | 7.52e-09
66 | 0.561 | 0.142 | 74.554
70 | 0.517 | 0.13 | 74.558
74 | 0.455 | 0.115 | 74.565
78 | 0.38 | 0.097 | 74.569
82 | 0.306 | 0.077 | 74.576
86 | 0.223 | 0.055 | 74.584
90 | 0.147 | 0.036 | 74.591
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.3282 |  | 138.602

## d2
Cavity No | Sample | Comment
1 | EP30-2 Quartz Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-09-27 | 9.5 | 5 | 129.5 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
25.1 | 7,4E-10 | 25 | 44.7504
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.557 | 0.152 | 74.551 | 2.4079e-05 | 1.23e-08
270 | 0.53 | 0.144 | 74.554
274 | 0.497 | 0.134 | 74.557
278 | 0.452 | 0.117 | 74.563
282 | 0.388 | 0.098 | 74.569
286 | 0.3 | 0.077 | 74.578
290 | 0.226 | 0.056 | 74.584
294 | 0.144 | 0.034 | 74.591
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.2965 |  | 138.4337

## d3
Cavity No | Sample | Comment
1 | EP30-2 Quartz Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-09-29 | 9 | 7 | 177 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
24.9 | 7.4e-10 | 25.1 | 44.7504
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.631 | 0.151 | 74.55 | 2.4127e-05 | 1.39e-08
270 | 0.595 | 0.144 | 74.553
274 | 0.534 | 0.133 | 74.557
278 | 0.472 | 0.117 | 74.562
282 | 0.398 | 0.099 | 74.569
286 | 0.309 | 0.077 | 74.578
290 | 0.223 | 0.056 | 74.584
294 | 0.147 | 0.035 | 74.591
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.3445 |  | 138.1574

## d4
Cavity No | Sample | Comment
1 | EP30-2 Quartz Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-10-01 | 10.5 | 9 | 226.5 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
24.3 | 3.5e-09 | 25 | 44.75
| Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.692 | 0.152 | 74.5525 | 2.4085e-05 | 1.77e-08
270 | 0.67 | 0.147 | 74.5532
274 | 0.62 | 0.136 | 74.5577
278 | 0.55 | 0.119 | 74.5637
282 | 0.47 | 0.101 | 74.569
286 | 0.39 | 0.078 | 74.5765
290 | 0.275 | 0.056 | 74.584
294 | 0.192 | 0.038 | 74.5907
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.2658 |  | 138.3977

## d5
Cavity No | Sample | Comment
1 | EP30-2 Quartz Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-10-04 | 11 | 12 | 299 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
26.2 | 2.9e-09 | 25 | 44.75
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.632 | 0.153 | 74.551 | 2.4009e-05 | 9.8e-09
270 | 0.61 | 0.147 | 74.554
274 | 0.571 | 0.136 | 74.557
278 | 0.51 | 0.12 | 74.56
282 | 0.436 | 0.101 | 74.5675
286 | 0.346 | 0.079 | 74.5765
290 | 0.257 | 0.057 | 74.5833
294 | 0.174 | 0.038 | 74.59
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.2768 |  | 138.8379

## d6
Cavity No | Sample | Comment
1 | EP30-2 Quartz Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-10-06 | 16 | 14 | 352 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
26.4 | 2.6e-09 | 25 | 44.7504
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.761 | 0.153 | 74.5525 | 2.4088e-05 | 2.84e-08
270 | 0.711 | 0.146 | 74.5533
274 | 0.68 | 0.135 | 74.5585
278 | 0.601 | 0.12 | 74.5638
282 | 0.51 | 0.1 | 74.5698
286 | 0.415 | 0.079 | 74.5773
290 | 0.3 | 0.057 | 74.584
294 | 0.2 | 0.038 | 74.5923
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.3018 |  | 138.3809

## d7
Cavity No | Sample | Comment
1 | EP30-2 Quartz Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-10-08 |  | 16 | 384 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
30.3 | 7.4e-10 | 26.9 | 44.7504
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.64 | 0.119 | 74.559 | 2.3282e-05 | 9.33e-09
270 | 0.628 | 0.114 | 74.561
274 | 0.58 | 0.105 | 74.566
278 | 0.515 | 0.093 | 74.57
282 | 0.44 | 0.078 | 74.575
286 | 0.358 | 0.061 | 74.583
290 | 0.264 | 0.044 | 74.59
294 | 0.174 | 0.028 | 74.598
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.5831 |  | 143.1699

## d8
Cavity No | Sample | Comment
1 | EP30-2 Quartz Disks | RESTARTED TESTING
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-10-14 | 9.5 | 22 | 537.5 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
25.1 | 1.9e-09 | 24.9 | 44.7509
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.381 | 0.111 | 74.56 | 2.4214e-05 | 8.42e-09
270 | 0.366 | 0.108 | 74.561
274 | 0.334 | 0.099 | 74.565
278 | 0.278 | 0.088 | 74.569
282 | 0.236 | 0.073 | 74.574
286 | 0.183 | 0.057 | 74.579
290 | 0.14 | 0.042 | 74.584
294 | 0.095 | 0.027 | 74.589
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.2969 |  | 137.6623

## d9
Cavity No | Sample | Comment
1 | EP30-2 Quartz Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-10-15 | 8.5 | 23 | 560.5 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
24.7 | 1.8e-09 | 24.9 | 44.7509
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.45 | 0.118 | 74.557 | 2.41e-05 | 1.11e-08
270 | 0.443 | 0.113 | 74.56
274 | 0.4 | 0.104 | 74.562
278 | 0.355 | 0.09 | 74.567
282 | 0.299 | 0.077 | 74.572
286 | 0.238 | 0.06 | 74.577
290 | 0.175 | 0.044 | 74.583
294 | 0.11 | 0.028 | 74.589
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.3134 |  | 138.3141

## d10
Cavity No | Sample | Comment
1 | EP30-2 Quartz Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-10-19 | 9 | 27 | 657 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
24.4 | 7.4e-10 | 25 | 44.7509
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.416 | 0.129 | 74.553 | 2.3963e-05 | 9.68e-09
270 | 0.388 | 0.124 | 74.554
274 | 0.359 | 0.113 | 74.558
278 | 0.31 | 0.1 | 74.564
282 | 0.271 | 0.084 | 74.569
286 | 0.215 | 0.067 | 74.574
290 | 0.162 | 0.048 | 74.58
294 | 0.104 | 0.031 | 74.587
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.3032 |  | 139.1013

## d11
Cavity No | Sample | Comment
1 | EP30-2 Quartz Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-10-20 | 10 | 28 | 682 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
24.3 | 7.4e-10 | 25 | 44.7509
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.443 | 0.132 | 74.552 | 2.4003e-05 | 1.03e-08
270 | 0.424 | 0.125 | 74.554
274 | 0.39 | 0.115 | 74.558
278 | 0.34 | 0.102 | 74.563
282 | 0.284 | 0.084 | 74.569
286 | 0.227 | 0.067 | 74.575
290 | 0.161 | 0.048 | 74.581
294 | 0.107 | 0.031 | 74.587
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.3157 |  | 138.8712

## d12
Cavity No | Sample | Comment
1 | EP30-2 Quartz Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-10-22 | 9 | 30 | 729 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
23.9 | 7.4e-10 | 24.9 | 44.7522
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.561 | 0.132 | 74.551 | 2.4122e-05 | 1.41e-08
270 | 0.549 | 0.126 | 74.554
274 | 0.508 | 0.118 | 74.557
278 | 0.45 | 0.102 | 74.562
282 | 0.369 | 0.086 | 74.567
286 | 0.295 | 0.069 | 74.574
290 | 0.214 | 0.049 | 74.58
294 | 0.14 | 0.031 | 74.586
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.2947 |  | 138.1874

## d13
Cavity No | Sample | Comment
1 | EP30-2 Quartz Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-10-25 | 10 | 33 | 802 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
24.2 | 1.4e-09 | 25 | 44.7522
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.51 | 0.136 | 74.551 | 2.4052e-05 | 8.67e-09
270 | 0.483 | 0.129 | 74.554
274 | 0.435 | 0.119 | 74.557
278 | 0.384 | 0.104 | 74.563
282 | 0.319 | 0.087 | 74.568
286 | 0.248 | 0.07 | 74.575
290 | 0.179 | 0.049 | 74.581
294 | 0.12 | 0.033 | 74.587
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.3105 |  | 138.5863

## d14
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-10-27 | 8.5 | 35 | 848.5 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
23.6 | 1.1e-09 | 24.9 | 44.7522
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.397 | 0.12 | 74.554 | 2.3963e-05 | 9.95e-09
270 | 0.35 | 0.116 | 74.556
274 | 0.331 | 0.104 | 74.56
278 | 0.311 | 0.094 | 74.563
282 | 0.255 | 0.078 | 74.569
286 | 0.202 | 0.062 | 74.575
290 | 0.149 | 0.045 | 74.581
294 | 0.098 | 0.028 | 74.587
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.3502 |  | 139.1008

## d15
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-10-29 | 10 | 37 | 898 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
23.8 | 1.1e-09 | 25 | 40.0049
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.387 | 0.105 | 74.56 | 2.4062e-05 | 9.8e-09
270 | 0.37 | 0.1 | 74.563
274 | 0.343 | 0.093 | 74.566
278 | 0.29 | 0.079 | 74.569
282 | 0.248 | 0.068 | 74.573
286 | 0.2 | 0.053 | 74.578
290 | 0.147 | 0.039 | 74.584
294 | 0.098 | 0.024 | 74.588
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.2815 |  | 138.5319

## d16
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-11-02 | 8 | 41 | 992 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
24.7 | 1.3e-09 | 25 | 44.7481
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.38 | 0.112 | 74.561 | 2.4041e-05 | 1.28e-08
270 | 0.358 | 0.107 | 74.563
274 | 0.329 | 0.099 | 74.566
278 | 0.287 | 0.087 | 74.57
282 | 0.248 | 0.072 | 74.576
286 | 0.195 | 0.057 | 74.581
290 | 0.15 | 0.042 | 74.587
294 | 0.099 | 0.028 | 74.591
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.3663 |  | 138.6509

## d17
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-11-03 | 8 | 42 | 1016 | 2.49 | Note: date & time were missing. I filled in a date intermediate between d16 & d18
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
22.5 | 8.7e-10 | 24.9 | 44.7481
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.282 | 0.091 | 74.564 | 2.3935e-05 | 7.63e-09
270 | 0.266 | 0.086 | 74.564
274 | 0.246 | 0.078 | 74.568
278 | 0.222 | 0.07 | 74.571
282 | 0.183 | 0.058 | 74.575
286 | 0.148 | 0.046 | 74.579
290 | 0.109 | 0.034 | 74.584
294 | 0.072 | 0.022 | 74.588
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.3592 |  | 139.2678

## d18
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-11-05 | 8 | 44 | 1064 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
21.6 | 7.7e-10 | 25 | 44.9812
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.286 | 0.086 | 74.563 | 2.3541e-05 | 1.24e-08
270 | 0.279 | 0.084 | 74.564
274 | 0.261 | 0.079 | 74.567
278 | 0.233 | 0.069 | 74.572
282 | 0.208 | 0.058 | 74.575
286 | 0.158 | 0.045 | 74.579
290 | 0.12 | 0.033 | 74.584
294 | 0.08 | 0.022 | 74.588
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.438 |  | 141.5985

## d19
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-11-08 | 8 | 47 | 1136 | 2.49 | Note: date & time were missing. I filled in a date intermediate between d18 & d20
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
23.1 | 9.1e-10 | 25 | 44.9212
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.615 | 0.147 | 74.544 | 2.3581e-05 | 1.91e-08
270 | 0.589 | 0.14 | 74.545
274 | 0.539 | 0.13 | 74.55
278 | 0.484 | 0.115 | 74.557
282 | 0.399 | 0.095 | 74.564
286 | 0.319 | 0.075 | 74.571
290 | 0.233 | 0.054 | 74.581
294 | 0.153 | 0.035 | 74.586
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.4696 |  | 141.3542

## d20
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-11-10 | 8.5 | 49 | 1184.5 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
21.9 | 7.5e-10 | 24.9 | 44.9212
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.683 | 0.143 | 74.542 | 2.3542e-05 | 9.14e-09
270 | 0.66 | 0.137 | 74.545
274 | 0.614 | 0.126 | 74.548
278 | 0.549 | 0.11 | 74.554
282 | 0.463 | 0.093 | 74.561
286 | 0.372 | 0.074 | 74.569
290 | 0.279 | 0.054 | 74.577
294 | 0.188 | 0.034 | 74.584
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.4703 |  | 141.5934

## d21
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-11-12 | 12 | 51 | 1236 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
23.8 | 7.4e-10 | 25 | 44.9212
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.859 | 0.151 | 74.539 | 2.3493e-05 | 1.14e-08
270 | 0.83 | 0.143 | 74.542
274 | 0.789 | 0.133 | 74.5465
278 | 0.683 | 0.117 | 74.554
282 | 0.577 | 0.098 | 74.5608
286 | 0.463 | 0.078 | 74.5675
290 | 0.332 | 0.056 | 74.5765
294 | 0.215 | 0.036 | 74.5833
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.4609 |  | 141.885

## d22
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-11-17 | 12 | 56 | 1356 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
23 | 7.5e-10 | 24.9 | 44.799
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.82 | 0.15 | 74.543 | 2.3685e-05 | 7.38e-09
270 | 0.777 | 0.143 | 74.546
274 | 0.723 | 0.132 | 74.551
278 | 0.636 | 0.116 | 74.557
282 | 0.52 | 0.097 | 74.564
286 | 0.402 | 0.078 | 74.572
290 | 0.296 | 0.055 | 74.58
294 | 0.199 | 0.035 | 74.586
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.4282 |  | 140.739

## d23
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-11-19 | 8.5 | 58 | 1400.5 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
23.9 | 1.1e-09 | 24.9 | 44.7664
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.705 | 0.15 | 74.547 | 2.3837e-05 | 7.84e-09
270 | 0.639 | 0.145 | 74.548
274 | 0.581 | 0.133 | 74.554
278 | 0.517 | 0.118 | 74.558
282 | 0.441 | 0.098 | 74.566
286 | 0.36 | 0.078 | 74.572
290 | 0.28 | 0.056 | 74.58
294 | 0.194 | 0.037 | 74.587
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.335 |  | 139.8396

## d24
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-11-22 | 10 | 61 | 1474 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
22.4 | 7.4e-10 | 24.9 | 44.7825
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.752 | 0.152 | 74.536 | 2.3426e-05 | 1.43e-08
270 | 0.737 | 0.145 | 74.539
274 | 0.691 | 0.132 | 74.545
278 | 0.604 | 0.118 | 74.55
282 | 0.52 | 0.098 | 74.56
286 | 0.417 | 0.077 | 74.566
290 | 0.306 | 0.057 | 74.577
294 | 0.206 | 0.038 | 74.583
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.5721 |  | 142.2918

## d25
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-11-24 | 8.5 | 63 | 1520.5 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
22.5 | 7.4e-10 | 24.9 | 44.7825
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.642 | 0.154 | 74.536 | 2.3395e-05 | 9.03e-09
270 | 0.615 | 0.147 | 74.541
274 | 0.582 | 0.134 | 74.545
278 | 0.509 | 0.118 | 74.551
282 | 0.43 | 0.099 | 74.557
286 | 0.34 | 0.079 | 74.566
290 | 0.253 | 0.058 | 74.574
294 | 0.166 | 0.036 | 74.583
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.4628 |  | 142.4779

## d26
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-11-29 | 9 | 68 | 1641 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
22.2 | 7.4e-10 | 24.9 | 44.782
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.54 | 0.142 | 74.541 | 2.3395e-05 | 8.65e-09
270 | 0.523 | 0.135 | 74.543
274 | 0.479 | 0.125 | 74.548
278 | 0.435 | 0.109 | 74.553
282 | 0.361 | 0.092 | 74.56
286 | 0.28 | 0.073 | 74.569
290 | 0.213 | 0.052 | 74.577
294 | 0.141 | 0.035 | 74.584
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.5343 |  | 142.482

## d27
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-12-01 | 9 | 70 | 1689 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
22.3 | 7.4e-10 | 25 | 44.7847
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.75 | 0.153 | 74.539 | 2.3415e-05 | 8.68e-09
270 | 0.725 | 0.146 | 74.54
274 | 0.69 | 0.134 | 74.546
278 | 617 | 0.12 | 74.552
282 | 0.512 | 0.098 | 74.56
286 | 0.421 | 0.08 | 74.567
290 | 0.314 | 0.058 | 74.575
294 | 0.207 | 0.037 | 74.584
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.4779 |  | 142.3602

## d28
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-12-03 | 8 | 72 | 1736 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
21.4 | 7.4e-10 | 24.9 | 44.7847
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.929 | 0.153 | 74.538 | 2.3424e-05 | 1.58e-08
270 | 0.902 | 0.147 | 74.541
274 | 0.84 | 0.136 | 74.545
278 | 0.755 | 0.119 | 74.551
282 | 0.614 | 0.099 | 74.56
286 | 0.502 | 0.08 | 74.566
290 | 0.369 | 0.059 | 74.575
294 | 0.236 | 0.037 | 74.582
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.4424 |  | 142.3027

## d29
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-12-06 | 14 | 75 | 1814 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
21.9 | 7.4e-10 | 24.9 | 44.7847
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.723 | 0.146 | 74.544 | 2.3535e-05 | 1.3e-08
270 | 0.694 | 0.141 | 74.545
274 | 0.64 | 0.129 | 74.553
278 | 0.564 | 0.113 | 74.556
282 | 0.475 | 0.095 | 74.563
286 | 0.382 | 0.074 | 74.571
290 | 0.281 | 0.055 | 74.58
294 | 0.181 | 0.035 | 74.587
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.4611 |  | 141.6318

## d30
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-12-08 | 9 | 77 | 1857 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
21.5 | 7.4e-10 | 24.9 | 447861
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.924 | 0.15 | 74.539 | 2.3461e-05 | 8.34e-09
270 | 0.903 | 0.143 | 74.542
274 | 0.854 | 0.133 | 74.546
278 | 0.749 | 0.117 | 74.554
282 | 0.644 | 0.099 | 74.559
286 | 0.516 | 0.079 | 74.569
290 | 0.366 | 0.055 | 74.577
294 | 0.244 | 0.036 | 74.585
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.5192 |  | 142.0781

## d31
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-12-10 | 10 | 79 | 1906 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
21.4 | 7.4e-10 | 24.9 | 44.7806
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 0.997 | 0.15 | 74.54 | 2.3483e-05 | 8.26e-09
270 | 0.962 | 0.143 | 74.543
274 | 0.881 | 0.13 | 74.548
278 | 0.781 | 0.116 | 74.552
282 | 0.652 | 0.096 | 74.563
286 | 0.52 | 0.077 | 74.569
290 | 0.382 | 0.056 | 74.577
294 | 0.245 | 0.036 | 74.584
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.4784 |  | 141.9473

## d32
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-12-14 | 9.5 | 83 | 2001.5 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
21.4 | 7.4e-10 | 24.9 | 44.7806
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.031 | 0.148 | 74.542 | 2.3362e-05 | 9.07e-09
270 | 0.988 | 0.143 | 74.543
274 | 0.901 | 0.131 | 74.548
278 | 0.802 | 0.116 | 74.554
282 | 0.67 | 0.097 | 74.562
286 | 0.53 | 0.077 | 74.57
290 | 0.391 | 0.056 | 74.578
294 | 0.255 | 0.036 | 74.584
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.4666 |  | 142.6847

## d33
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-12-15 | 15 | 84 | 2031 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
21.3 | 7.4e-10 | 24.9 | 44.7806
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.035 | 0.146 | 74.543 | 2.3395e-05 | 9.66e-09
270 | 0.975 | 0.141 | 74.545
274 | 0.894 | 0.129 | 74.549
278 | 0.78 | 0.114 | 74.556
282 | 0.653 | 0.096 | 74.562
286 | 0.517 | 0.077 | 74.57
290 | 0.38 | 0.056 | 74.578
294 | 0.245 | 0.036 | 74.585
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.4583 |  | 142.4781

## d34
Cavity No | Sample | Comment
1 | EP30-2 Bonded Quartz Glass Disks
Loading Date | Measurement Date | Mea. Time (24h) | Tested Days | Tested Hours | Power Monitor ratio
2010-09-22 | 2010-12-17 | 14 | 86 | 2078 | 2.49
Room T (°C) | Vacuum (Torr) | Cavity T (°C) | LT (°C)
21.4 | 7.4e-10 | 24.9 | 44.7806
λ/2 WP position | Power In Monitor (V) | Power Out Monitor (W) | TEM00-01 Beat Freq (MHz) | Ringdown Time (s) | Ringdown fitting σ (s)
264 | 1.031 | 0.146 | 74.544 | 2.3462e-05 | 8.49e-09
270 | 0.979 | 0.138 | 74.545
274 | 0.901 | 0.128 | 74.549
278 | 0.783 | 0.112 | 74.556
282 | 0.653 | 0.094 | 74.562
286 | 0.52 | 0.075 | 74.57
290 | 0.368 | 0.054 | 74.578
294 | 0.239 | 0.036 | 74.585
|  | Fitted Two Mirrors' Total absorption (ppm) |  | Calculated Round Trip Total Loss (ppm)
|  | 1.4476 |  | 142.0712