| APPROVALS | DATE | REV | DCN NO. | BY | CHECK | DCC | DATE |
| --- | --- | --- | --- | --- | --- | --- | --- |
| AUTHOR: G. Billingsley | 7-29-03 | A | E030244-00 |  |  |  | 7-29-03 |
| CHECKED: |  |  |  |  |  |  |  |
| APPROVED: |  |  |  |  |  |  |  |

Purpose:  These Flat-Flat sapphire substrates are to be used in investigations into sapphire Q and thermal noise.

Applicable Documents:  LIGO-D030403 70 x 60 mm sapphire Q substrate

Material:  HEMLITE or Better single crystal sapphire

Orientation:  Different cylinder axis for each of the three types:  a, c and m ± 2 degrees

Diameter:  70.0 mm +/- 0.2 mm

Thickness:  60.0 mm +/- 0.2 mm

Wedge:  30 ± 5 minutes, parallel to the c axis of the material within 3 degrees, except for the c-plane piece, where the wedge should be parallel to the a axis of the material within 3 degrees

Other Mechanical dimensions:  Bevel, cylindricity and runout to best commercial practice

Polish:  All surfaces polished out of gray, including bevel and cylinder

Surface Figure:  Nominally flat, /4, PV at 633 nm

Microroughness:  Super polish, front side only, < 2 Angstroms RMS