LIGO Laboratory / LIGO Scientific Collaboration

LIGO-T040032-01-D	Advanced LIGO	25 April 2006

Polyimide Hydrocarbon Outgas Rate
and Temperature Limits for the Isolation System Actuator

Dennis Coyne

Distribution of this document:

LIGO Science Collaboration

This is an internal working note

of the LIGO Project.

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW17-161 175 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Introduction

The hydrocarbon outgas rate for polyimide materials is derived from existing mass spectrum test data. The outgas rate is then used to set a maximum temperature limit on the voice coil actuator for the advanced LIGO Seismic Isolation System (SEI).

# Polyimide Outgas Rate

In order to set a maximum in-vacuum operating temperature for the voice coil, we would ideally have measurements of the outgas spectrum and rate, including temperature dependence, for the intended SEI voice coil actuator (or at least it's materials). The voice coil actuators used to date (BEI/Kimco LA50-62-004Z and LA18-32-006Z) have not been developed for UHV use. It is likely that the epoxies, wire insulation and potting compound outgas at too high of a rate (though this has not been confirmed by measurement). A version of the smaller actuator (LA18-32-006Z) with kapton insulated wire potted with Cycom 3001 (a high temperature polyimide adhesive) has been produced. The mass spectrum of this unit, after cleaning and vacuum baking at 200C for 48 hours has been measured. The hydrocarbon (HC) outgassing rate derived from this residual gas assay (RGA) is suspect due to lack of repeatability of measurements on this chamber. The HC rate (sum of AMUs 41, 43, 53, 55 and 57) is 3.8E-10 torr-liter/sec (very likely limited by the background outgassing rate of the chamber). In addition to being suspect, this measurement was only made a room temperature. In order to establish a maximum in-vacuum operating temperature, we must know the doubling temperature (or some other measure of the outgas rate dependence with temperature). A repeat vacuum bake and mass spectrum measurement of the voice coil is planned in March, 2004 in the LHO vacuum bake oven with measurements at ~70C, ~45C and room temperature (~20 C).

In the interim we might use the outgas rate from other measurements of Polyimide materials to set approximate limits on the temperature of the voice coil. The mass spectrum of a number of polyimide materials have been measured. In particular, the Virgo measurements include the mass spectrum at multiple elevated temperatures (a practice that LIGO should adopt). Unfortunately it seems the Virgo group has not performed mass spectrum measurements of Cycom 3001 (confirmation has been requested). Five polyimide measurements were examined:

1) Coils with Polyamide-Polyimide Insulation and Polyimide Resin "varnish". The "varnish" is BASG 48816 polyimide/carbon pigment resin from Ball Aerospace.

2) Coils with Polyamide-Polyimide Insulation (without the varnish)

3) Cabling with Pyre-ML Insulation. Pyre-ML is a Dupont aromatic polyimide.

4) Kapton ribbon cables (Kapton is a Dupont trademark polyimide). Virgo cables were provide by Axon.

5) Kapton Ribbon cables with PEEK weave and PEEK connectors from MDC Vacuum Products Inc.

The Virgo report gives the total outgassing rate and the mass spectrum (in uncalibrated units) at multiple temperatures. The sequence was to dwell at each temperature and then take a mass spectrum measurement before increasing the temperature. The dwell time at temperature was not always long enough to reach a steady-state outgassing rate. The temperatures were typically 33 C, 50 C, 100 C, 150 C and then back to 40 C to 33 C. The mass spectrum plots were digitized and integrated to get a scale factor to calibrate the mass spectrum. This scale factor was then used with the HC flag AMU sum to get the HC outgassing rate.

The outgassing rate, k, is assumed to follow an Arrhenius rate equation:

where in general the pre-expenential factor, A, and the activation energy, E, are temperature dependent; Here these coefficients are assumed to be temperature independent. The temperature change required to double the outgassing rate is:

Of course a single rate equation may not apply. In addition measurements at significantly higher temperatures than the operating temperature range may bring different kinetics into play. However the required hydrocarbon outgas rate is too low to measure at room temperature.

## Polyimide Resin

For the Polyimide Resin, the outgassing rate (torr-liter/sec/cm2) versus temperature is given in Figure 1. The same data is plotted versus 1/T in Figure 2. The outgassing rate at the two low temperatures (33 C pre-bake and 40 C post-bake) is limited by the mass spectrometer noise floor (or at least by the plot scale in the Virgo report). Also shown is a fit to an Arrhenius equation.

Figure 1: Polyimide Resin Outgas Rate vs Temperature

Figure 2: Polyimide Resin Outgas Rate vs Temperature

A least squares fit to the outgas rate at the three temperatures high enough to measure the outgas rate not limited by the noise floor of the spectrometer (or the lower limit of the plot in the Virgo report), yields:

A = 3.09e-7 torr-liter/sec/cm2

E = 268 C

T = 1.03 C

## Polyimide Insulation

Figure 3: Polyimide Insulation Outgas Rate vs Temperature

A = 1.03e-8 torr-liter/sec/cm2

E = 286 C

T = 0.97 C

## Pyre-ML Insulation

The Pyre-ML data is limited by the noise floor (limit of the plot in the Virgo report).

Figure 4: Pyre-ML Insulation Outgas Rate vs Temperature

A = 5.46e-10 torr-liter/sec/cm2

E = 245 C

T = 1.13 C

## Axon Kapton Cabling

The Virgo report on Kapton cabling from Axon is still to be further analyzed. Using the 80 C and 150 C data (post-bake), one gets the following Arrhenius constants:

A = 1.35e-9 torr-liter/sec/cm2

E = 753 C

T = 0.37 C

With the admittedly uncertain fit, the room temperature (20 C) HC outgas rate is 6e-26 torr-liter/sec/cm2. If the activation energy from the previous fits is used (E ~ 280 C), then the room temperature HC outgas rate is 1e-15 torr-liter/sec/cm2, which is consistent with the measured MDC kapton cable outgas rate.

## MDC Kapton Cabling

The HC outgas rate reported in E980274-00 is 2.80e-15 torr-liter/sec/cm2 (each cable has 6080 cm2 of area). This value does not have the oven or RGA background subtracted. Since empty chamber measurements were extremely infrequent in the Caltech vacuum bake ovens used for initial LIGO bake loads, we do not have a background to subtract. In addition, the RGA was not capable of being valved off from the chamber, so an RGA background can not be subtracted either. Looking at clean load scans (metal) in the same chamber around the same time frame as the E980274-00 load, it seems that the oven HC background is about at ½ the measured level, so a reasonable estimate of the kapton outgas rate is 1.4e-15 torr-liter/sec/cm2.

# Voice Coil Maximum Temperature

The exposed surface area of the larger voice coil is estimated to be 143 cm2. With the MDC Kapton outgas level reported above, this means at room temperature (20C), the expected HC outgas rate is 2e-13 torr-liter/sec/coil.

The allowed outgas rate for the SEI voice coil is 3.8e-12 torr-liter/sec/coil. Using an activation energy of 245 (the lowest reported above), the allowable temperature is only 26 C, or 6 C above ambient.

# Appendix: Cycom 3002 Datasheet

# Footnotes

A datasheet for Cycom 3001 has been included in the appendix of this memo. A datasheet for Cycom 3002 (a prepeg with Cycom 3001 structural resin) has also been included in the appendix of this memo.

Virgo document VIR-TRE-PIS-3400-114, "Outgassing Test of Coils with Polyamide-polyimide Insulation and Varnishing".

Virgo document VIR-TRE-PIS-3400-113, "Outgassing Test of Coils with Polyamide-polyimide Insulation".

Virgo document VIR-TRE-PIS-3400-115, "Outgassing Test of Cabling with Pyre-ML Insulation".

Virgo document VIR-TRE-PIS-3400-137, "Outgassing Test of an Axon Kapton Ribbon".

Vacuum Bake Traveler for Kapton Ribbon Cable, LIGO-E980274-00. This vacuum bake load appears to have the lowest outgassing rate measured at Caltech for initial LIGO kapton cables. Vacuum bake load travelers from LHO should be checked as well.

D. Coyne, Vacuum Hydrocarbon Outgassing Requirements", T040001-00.