*LIGO Laboratory / LIGO Scientific Collaboration*

LIGO-T1100160-v8 *Advanced LIGO* August 29th, 2012

**aLIGO ISC QPD Transimpedance Amplifier Test Procedure**

Mohana Mageswaran, R. Abbott

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

|  |  |
| --- | --- |
| **California Institute of Technology**  **LIGO Project – MS 18-34**  **1200 E. California Blvd.**  **Pasadena, CA 91125**  Phone (626) 395-2129  Fax (626) 304-9834  E-mail: info@ligo.caltech.edu | **Massachusetts Institute of Technology**  **LIGO Project – NW22-295**  **185 Albany St**  **Cambridge, MA 02139**  Phone (617) 253-4824  Fax (617) 253-7014  E-mail: info@ligo.mit.edu |
| **LIGO Hanford Observatory**  **P.O. Box 1970**  **Richland WA 99352**  Phone 509-372-8106  Fax 509-372-8137 | **LIGO Livingston Observatory**  **P.O. Box 940**  **Livingston, LA 70754**  Phone 225-686-3100  Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Overview

This test procedure applies to ISC QPD Transimpedance Amplifier circuit board LIGO-D1001974-v2 contained within chassis assembly D1002481. A block diagram of the ISC QPD Transimpedance circuit board is shown in Figure 1. Two such QPD Transimpedance Amps and one ISC QPD Transimpedance Amplifier Interface are packaged in one chassis.



Figure 1 QPD Transimpedance Amplifier Circuit Block Diagram

# Testing

Each production chassis must be functionally tested and the results recorded in Section 4. It is assumed that the person using this procedure is familiar with Dynamic Signal Analyzers, and rudimentary test equipment including oscilloscopes and multimeters.

**Serial Number Data**

* Record all serial number data in Table 1

**DC Tests**

Apply +/- 18, +/-200 mV Volts DC to the chassis under test and record front panel LED operation, total positive and negative power supply current, internal regulator output voltage and individual circuit board power supply currents as required in

* Table 2.

# Reference for chassis front and rear panel layout

Figure 2: QPD Transimpedance Amplifier Chassis Front Panel



Figure 3: QPD Transimpedance Amplifier Chassis Rear Panel



# Test Data Tables

## General Information

|  |  |  |  |
| --- | --- | --- | --- |
| **Tested By** |  | **Date** |  |

Table 1 Serial Number Data

|  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| **Chassis Serial Number** | **DC PWR Board**  **PCB Serial #** | **Amplifier 1**  **PCB Serial #** | **Amplifier 2**  **PCB Serial #** | **Interface Board**  **PCB Serial #** | |  |  | |
|  |  |  |  | |  | | |

## DC Power Supply Data

Total chassis and individual circuit board quiescent current draw is recorded in

Table 2. For the individual circuit boards, unplug all but one board at a time and record the chassis current draw of the +/- 18VDC supply. Use caution in believing the digital readouts of laboratory triple output power supplies. Their meters are not highly accurate. When in doubt, use a multimeter on the appropriate scale in series with the supply to be measured.

Table 2, Record of DC Test Data

|  |  |  |  |
| --- | --- | --- | --- |
| **Parameter** | **Typical Value** | **Allowable Range** | **Measured Value** |
| Front Panel +/- 15VDC Power LEDs | Both Lit | N/A |  |
| Rear Panel +/- 15VDC  Power LEDs | Both Lit | N/A |  |
| +18VDC, +/-0.2VDC **TOTAL** supply current | 270 mA | +/- 50mA |  |
| -18VDC, +/-0.2VDC **TOTAL** supply current | 240 mA | +/- 50mA |  |
| Regulated Internal DC Voltage under full load (both boards) | 15 VDC | +/- 0.5VDC |  |
| Regulated Internal DC Voltage under full load (both boards) | -15 VDC | +/- 0.5VDC |  |

## DC Offsets on Each Differential Output

As a general measure of the health, the DC offset at the differential outputs for each channel must be measured. Using a multimeter, measure the DC offset at each differential output on the associated rear panel D-sub connector. The input connector is to be left open. Record the results in Table 3.

Table 3, Differential Output DC Offset

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| ***Differential DC Measurement Point*** | ***Typical DC Offset*** | ***Allowable Range*** | ***Measured DC Offset*** | |
| **Amplifier 1** | **Amplifier 2** |
| Channel 1 | 0VDC | +/- 5mV |  |  |
| Channel 2 | 0VDC | +/- 5mV |  |  |
| Channel 3 | 0VDC | +/- 5mV |  |  |
| Channel 4 | 0VDC | +/- 5mV |  |  |

**4.4 Transimpedance**

Calculate the transimpedance by using the laboratory Voltage/Current calibrator Model IVC-222HP 11.

Inject 1mA DC into the appropriate anode input of the QPD. Measure the DC Voltage at the differential output; calculate the trans impedance of the circuit by:

***Transimpedance = Vout / 1mA***

Table 4, Differential Output Transimpedance

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| ***Differential DC Measurement Point*** | ***Transimpedance*** | ***Allowable Range*** | ***Calculated Transimpedance*** | |
| **Amplifier 1** | **Amplifier 2** |
| Channel 1 | 1K | +/- 5mΩ |  |  |
| Channel 2 | 1K | +/- 5mΩ |  |  |
| Channel 3 | 1K | +/- 5mΩ |  |  |
| Channel 4 | 1K | +/- 5mΩ |  |  |

## Frequency Response

The transfer function of each channel of the amplifier should be measured using an SR785 dynamic signal analyzer. The input impedance to all channels of this circuit is 10 ohms. Due to this low impedance, a 1kΩ resistor is required to be placed in series with the SR785 source. A simple set of clip leads and a breakout board is sufficient. The SR785 input drive level is 10mV for all swept sine measurements.

Measure the magnitude and the phase differentially at the rear panel D-sub output for each channel as required. Record the results the following tables.

Table 5, Noise Cancellation Amp

|  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| **Measurement**  **Frequency** | **Magnitude (dB)** | **Allowable Range** | **Phase (deg)** | **Allowable Range** | **Measured Magnitude** | **Measured Phase** | **Pass/Fail** |
| 10Hz | 34 | +/- 1dB | 73.6 | +/- 5 deg |  |  |  |
| 100Hz | 45 | +/- 1dB | 21.4 | +/- 5 deg |  |  |  |
| 1KHz | 46 | +/- 1dB | 0.9 | +/- 5 deg |  |  |  |
| 10KHz | 46 | +/- 1dB | -12.8 | +/- 5 deg |  |  |  |
| 100KHz | 38 | +/- 1dB | -76.1 | +/- 5 deg |  |  |  |

Table 6, Frequency Response Amp 1\_Quadrant 1

|  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| **Measurement**  **Frequency** | **Magnitude (dB)** | **Allowable Range** | **Phase (deg)** | **Allowable Range** | **Measured Magnitude** | **Measured Phase** | **Pass/Fail** |
| 10Hz | 34 | +/- 1dB | -106 | +/- 5 deg |  |  |  |
| 100Hz | 45 | +/- 1dB | -159 | +/- 5 deg |  |  |  |
| 1KHz | 46 | +/- 1dB | -179 | +/- 5 deg |  |  |  |
| 10KHz | 46 | +/- 1dB | 168 | +/- 5 deg |  |  |  |
| 100KHz | 38 | +/- 1dB | 108 | +/- 5 deg |  |  |  |

Table 7, Frequency Response Amp 1\_Quadrant 2

|  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| **Measurement**  **Frequency** | **Magnitude (dB)** | **Allowable Range** | **Phase (deg)** | **Allowable Range** | **Measured Magnitude** | **Measured Phase** | **Pass/Fail** |
| 10Hz | 34 | +/- 1dB | -106 | +/- 5 deg |  |  |  |
| 100Hz | 45 | +/- 1dB | -159 | +/- 5 deg |  |  |  |
| 1KHz | 46 | +/- 1dB | -179 | +/- 5 deg |  |  |  |
| 10KHz | 46 | +/- 1dB | 168 | +/- 5 deg |  |  |  |
| 100KHz | 38 | +/- 1dB | 108 | +/- 5 deg |  |  |  |

Table 8, Frequency Response Amp 1\_Quadrant 3

|  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| **Measurement**  **Frequency** | **Magnitude (dB)** | **Allowable Range** | **Phase (deg)** | **Allowable Range** | **Measured Magnitude** | **Measured Phase** | **Pass/Fail** |
| 10Hz | 34 | +/- 1dB | -106 | +/- 5 deg |  |  |  |
| 100Hz | 45 | +/- 1dB | -159 | +/- 5 deg |  |  |  |
| 1KHz | 46 | +/- 1dB | -179 | +/- 5 deg |  |  |  |
| 10KHz | 46 | +/- 1dB | 168 | +/- 5 deg |  |  |  |
| 100KHz | 38 | +/- 1dB | 108 | +/- 5 deg |  |  |  |

Table 9, Frequency Response Amp 1\_Quadrant 4

|  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| **Measurement**  **Frequency** | **Magnitude (dB)** | **Allowable Range** | **Phase (deg)** | **Allowable Range** | **Measured Magnitude** | **Measured Phase** | **Pass/Fail** |
| 10Hz | 34 | +/- 1dB | -106 | +/- 5 deg |  |  |  |
| 100Hz | 45 | +/- 1dB | -159 | +/- 5 deg |  |  |  |
| 1KHz | 46 | +/- 1dB | -179 | +/- 5 deg |  |  |  |
| 10KHz | 46 | +/- 1dB | 168 | +/- 5 deg |  |  |  |
| 100KHz | 38 | +/- 1dB | 108 | +/- 5 deg |  |  |  |

Table 10, Frequency Response Amp2\_Quadrant 1

|  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| **Measurement**  **Frequency** | **Magnitude (dB)** | **Allowable Range** | **Phase (deg)** | **Allowable Range** | **Measured Magnitude** | **Measured Phase** | **Pass/Fail** |
| 10Hz | 34 | +/- 1dB | -106 | +/- 5 deg |  |  |  |
| 100Hz | 45 | +/- 1dB | -159 | +/- 5 deg |  |  |  |
| 1KHz | 46 | +/- 1dB | -179 | +/- 5 deg |  |  |  |
| 10KHz | 46 | +/- 1dB | 168 | +/- 5 deg |  |  |  |
| 100KHz | 38 | +/- 1dB | 108 | +/- 5 deg |  |  |  |

Table 11, Frequency Response Amp 2\_Quadrant 2

|  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| **Measurement**  **Frequency** | **Magnitude (dB)** | **Allowable Range** | **Phase (deg)** | **Allowable Range** | **Measured Magnitude** | **Measured Phase** | **Pass/Fail** |
| 10Hz | 34 | +/- 1dB | -106 | +/- 5 deg |  |  |  |
| 100Hz | 45 | +/- 1dB | -159 | +/- 5 deg |  |  |  |
| 1KHz | 46 | +/- 1dB | -179 | +/- 5 deg |  |  |  |
| 10KHz | 46 | +/- 1dB | 168 | +/- 5 deg |  |  |  |
| 100KHz | 38 | +/- 1dB | 108 | +/- 5 deg |  |  |  |

Table 12, Frequency Response Amp2\_Quadrant 3

|  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| **Measurement**  **Frequency** | **Magnitude (dB)** | **Allowable Range** | **Phase (deg)** | **Allowable Range** | **Measured Magnitude** | **Measured Phase** | **Pass/Fail** |
| 10Hz | 34 | +/- 1dB | -106 | +/- 5 deg |  |  |  |
| 100Hz | 45 | +/- 1dB | -159 | +/- 5 deg |  |  |  |
| 1KHz | 46 | +/- 1dB | -179 | +/- 5 deg |  |  |  |
| 10KHz | 46 | +/- 1dB | 168 | +/- 5 deg |  |  |  |
| 100KHz | 38 | +/- 1dB | 108 | +/- 5 deg |  |  |  |

Table 13, Frequency Response Amp 2\_Quadrant 4

|  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| **Measurement**  **Frequency** | **Magnitude (dB)** | **Allowable Range** | **Phase (deg)** | **Allowable Range** | **Measured Magnitude** | **Measured Phase** | **Pass/Fail** |
| 10Hz | 34 | +/- 1dB | -106 | +/- 5 deg |  |  |  |
| 100Hz | 45 | +/- 1dB | -159 | +/- 5 deg |  |  |  |
| 1KHz | 46 | +/- 1dB | -179 | +/- 5 deg |  |  |  |
| 10KHz | 46 | +/- 1dB | 168 | +/- 5 deg |  |  |  |
| 100KHz | 38 | +/- 1dB | 108 | +/- 5 deg |  |  |  |

Table 14, Noise Cancellation Amp 2

|  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| **Measurement**  **Frequency** | **Magnitude (dB)** | **Allowable Range** | **Phase (deg)** | **Allowable Range** | **Measured Magnitude** | **Measured Phase** | **Pass/Fail** |
| 10Hz | 34 | +/- 1dB | 73.6 | +/- 5 deg |  |  |  |
| 100Hz | 45 | +/- 1dB | 21.4 | +/- 5 deg |  |  |  |
| 1KHz | 46 | +/- 1dB | 0.9 | +/- 5 deg |  |  |  |
| 10KHz | 46 | +/- 1dB | -12.8 | +/- 5 deg |  |  |  |
| 100KHz | 38 | +/- 1dB | -76.1 | +/- 5 deg |  |  |  |

## Output Noise Spectra

The output noise voltage of each channel of the amplifier should be measured using the dynamic signal analyzer SR785. This measurement should be made while the input is open, and the frequency range is set from 1Hz to 100 KHz.

Measure the output referred noise differentially at the rear panel D-sub output for each channel as required. Record the results in Table to

Table

Table 15, Amp 1\_Quadrant 1 Noise

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Measurement Frequency** | **Typical Amplitude dBVrms/√Hz** | **Allowable Range** | **Measured Amplitude dBVrms/√Hz** | **Pass/Fail** |
| 10Hz | -129 | +/- 3dB |  |  |
| 100Hz | -118 | +/- 3dB |  |  |
| 1KHz | -118 | +/- 3dB |  |  |
| 10KHz | -118 | +/- 3dB |  |  |

Table 16, Amp 1\_Quadrant 2 Noise

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Measurement Frequency** | **Typical Amplitude dBVrms/√Hz** | **Allowable Range** | **Measured Amplitude dBVrms/√Hz** | **Pass/Fail** |
| 10Hz | -129 | +/- 3dB |  |  |
| 100Hz | -118 | +/- 3dB |  |  |
| 1KHz | -118 | +/- 3dB |  |  |
| 10KHz | -118 | +/- 3dB |  |  |

Table17, Amp 1\_Quadrant 3 Noise

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Measurement Frequency** | **Typical Amplitude dBVrms/√Hz** | **Allowable Range** | **Measured Amplitude dBVrms/√Hz** | **Pass/Fail** |
| 10Hz | -129 | +/- 3dB |  |  |
| 100Hz | -118 | +/- 3dB |  |  |
| 1KHz | -118 | +/- 3dB |  |  |
| 10KHz | -118 | +/- 3dB |  |  |

Table 18, Amp 1\_Quadrant 4 Noise

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Measurement Frequency** | **Typical Amplitude dBVrms/√Hz** | **Allowable Range** | **Measured Amplitude dBVrms/√Hz** | **Pass/Fail** |
| 10Hz | -129 | +/- 3dB |  |  |
| 100Hz | -118 | +/- 3dB |  |  |
| 1KHz | -118 | +/- 3dB |  |  |
| 10KHz | -118 | +/- 3dB |  |  |

Table 19, Amp 2\_Quadrant 1 Noise

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Measurement Frequency** | **Typical Amplitude dBVrms/√Hz** | **Allowable Range** | **Measured Amplitude dBVrms/√Hz** | **Pass/Fail** |
| 10Hz | -129 | +/- 3dB |  |  |
| 100Hz | -118 | +/- 3dB |  |  |
| 1KHz | -118 | +/- 3dB |  |  |
| 10KHz | -118 | +/- 3dB |  |  |

**Table 20, Amp2\_Quadrant 2 Noise**

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Measurement Frequency** | **Typical Amplitude dBVrms/√Hz** | **Allowable Range** | **Measured Amplitude dBVrms/√Hz** | **Pass/Fail** |
| 10Hz | -129 | +/- 3dB |  |  |
| 100Hz | -118 | +/- 3dB |  |  |
| 1KHz | -118 | +/- 3dB |  |  |
| 10KHz | -118 | +/- 3dB |  |  |

Table 21, Amp2\_Quadrant 3 Noise

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Measurement Frequency** | **Typical Amplitude dBVrms/√Hz** | **Allowable Range** | **Measured Amplitude dBVrms/√Hz** | **Pass/Fail** |
| 10Hz | -129 | +/- 3dB |  |  |
| 100Hz | -118 | +/- 3dB |  |  |
| 1KHz | -118 | +/- 3dB |  |  |
| 10KHz | -118 | +/- 3dB |  |  |

Table 22, Amp2\_Quadrant 4 Noise

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Measurement Frequency** | **Typical Amplitude dBVrms/√Hz** | **Allowable Range** | **Measured Amplitude dBVrms/√Hz** | **Pass/Fail** |
| 10Hz | -129 | +/- 3dB |  |  |
| 100Hz | -118 | +/- 3dB |  |  |
| 1KHz | -118 | +/- 3dB |  |  |
| 10KHz | -118 | +/- 3dB |  |  |