LIGO Laboratory / LIGO Scientific Collaboration

LIGO- T1100191-v1	Advanced LIGO UK      March 2011

Transmon Driver Unit Users’ Guide

R. M. Cutler, University of Birmingham

Distribution of this document:

Inform aligo_sus

This is an internal working note

of the Advanced LIGO Project, prepared by members of the UK team.

| Institute for Gravitational Research University of Glasgow Phone +44 (0) 141 330 5884 Fax +44 (0) 141 330 6833 E-mail k.strain@physics.gla.ac.uk | School of Physics and Astronomy University of Birmingham Phone +44 (0) 121 414 6447 Fax +44 (0) 121 414 3722 E-mail av@star.sr.bham.ac.uk |
| --- | --- |
| Engineering Department CCLRC Rutherford Appleton Laboratory Phone +44 (0) 1235 445 297 Fax +44 (0) 1235 445 843 E-mail J.Greenhalgh@rl.ac.uk | Department of Physics University of Strathclyde Phone +44 (0) 1411 548 3360 Fax +44 (0) 141 552 2891 E-mail N.Lockerbie@phys.strath.ac.uk |

http://www.ligo.caltech.edu/

http://www.physics.gla.ac.uk/igr/sus/

http://www.sr.bham.ac.uk/research/gravity/rh,d,2.html

http://www.eng-external.rl.ac.uk/advligo/papers_public/ALUK_Homepage.htm

Transmon Driver Users Guide

Contents

1. Document Organization

2. LIGO

2.1 The Mirror Suspension System

2.2    The Transmon Control System

3    The Transmon Drive Unit

3.1   The Drive card

3.2   The Monitor Card

4     Using the Unit

4.1   Connectors

4.2   Connector Functions

4.2.1  Connectors on the Front Panel

4.2.2  Connectors on the Rear Panel

4.3   Connector Details

5.   Functional description

5.1   Inside the Unit

5.2   The Drive Board

5.3   Monitor Board

6.  Assembly and Disassembly

## 1.  DOCUMENT ORGANIZATION

- Section 2 describes the Transmon control system, and the function of the OMC Drive Unit within the control system.

- Section 3 describes the Transmon Drive Unit

- Section 4 describes how to use the unit, and details how the Transmon Drive Unit connections to the rest of the system,

- Section 5 gives a functional description

- Section 6 describes assembly and disassembly

2 LIGO

The LIGO (Laser Interferometric Gravitational-wave Observatory) experiment is designed to detect gravity waves.

The laser interferometery method used involves mirrors which are held in a steady position to a very high degree of precision.

2.1 The Mirror Suspension System

Mirror stability is achieved by means of a suspension consisting heavy masses.

The purpose of the control system is to maintain the position of the Transmon. Velocity control also damps Transmon movements.

2.2 The Transmon Control System

The Transmon Driver forms part of the control system which controls the position and movement of the Transmon suspension electromagnetically. Fig 1 is a simplified diagram of the control system, illustrating the role of the Driver Unit.

Fig 1 The Transmon Control System

The demand signals to the Drive Amplifiers are generated by the Digital Control System, converted to analogue signals, and then filtered in the AI Unit. The signals are then amplified and filtered by the Drive Amplifiers and pass through the Satellite Box to provide the demanded currents to the drive coils. The position of the Transmon stage is indicated by mechanical “flags” to the OSEM optical position detectors. The position detectors consist of LED and photodiode pairs. The photodiode outputs are amplified in the Satellite Box, and then fed back to the Digital Control System via the filters in the AA unit to the A/D convertors.

3 THE TRANSMON DRIVE UNIT

Fig 2 The Transmon Driver unit

The Transmon Driver unit houses a Drive Card and a Monitor Card. Each of these cards has four channels.

3.1 The Drive Card

The Drive Card performs amplification and filtering of the demand signals. The first filter stage may be switched in or out by the relay control signals on Binary I/O 2, to give a selected trade off between noise and dynamic range. The state of the relays is monitored by Binary I/O 1.

The two filter stages in each filter channel act as low pass filters with a 1 Hz corner frequency. Each filter stage also has a “zero” at 10 Hz, so its attenuation is almost constant above 10 Hz.

3.2 The Monitor Card

The monitor card monitors the performance of the Driver, presenting the output Voltage, instantaneous Current and RMS current to the A/D circuits via the AA Unit. The output voltage noise from the amplifiers is also presented, after amplification, via a separate output connector.

4. USING THE UNIT

To use the unit, first connect it to the system as described below. The power switch should be in the “off” position at this stage (0). The input voltage should be checked before connection (Positive voltage = 17v nominal, 16.5 v minimum, 18v maximum, negative voltage = -17v nominal, -16.5 volts minimum, 18 volts maximum.) and A1 = positive, A2 = return, A3 = negative.

Having connected all the cables correctly and checked the power supply, the power may be switched on. The green LED indicator lights on the front and rear panels should then become illuminated, indicating that the on board regulators are operating.

4.1 Connectors

The input and output connections to the Transmon Drive Unit were shown on Figure 3. Details of these connectors follow.

4.2 Connector Functions

The unit needs to be connected to the system in the following way.

4.2.1 Connectors on the Front Panel

These will now be described starting from the left hand connector.

PD Out to AA socket

This needs to be connected to the AA unit. It transmits the photodiode outputs via the AA unit to the A/D convertors and control system.

LED Mon socket

This needs to be connected to the A/D converters via the AA unit. It consists of the signals which monitor the state of the LEDs in the satellite box.

Drive Input plug

This needs to be connected to the Coil demand signal from the D/A Converter, unit via the AI unit. The signals which are applied to this plug determine the coil drive current.

Test plug

This must normally be terminated by a shorting socket, with all connections joined together. It may also be used to feed demand signals in for test purposes.

Binary I/O 1 socket

The signals on this socket are used to monitor the status of the filter relays and the Test relay.

Binary I/O 2 plug

This connector is used to control the state of the relays which switch the filters in and out. It also controls the state of the relays which switch the Drive Amplifier inputs between the Drive Inputs and the Test Inputs.

4.2.2 Connectors on the Rear Panel

Starting from the left hand end of the rear panel:

Switch/Thermal cut out

This controls the DC power input to the unit. It incorporates a thermal cut out to protect against a prolonged over current situation.

DC In plug

This 3 pin heavy duty ‘D’ connector is to be connected to the power supply. The nominal input voltage is +/-17 volts.

Noise Monitor socket

This presents the outputs from the Noise Monitor circuits, and should be connected to the A/D converters via the AA unit.

V, I and RMS monitor plug

This presents the outputs from the Voltage Monitor, Current Monitor and R.M.S Monitor circuits, and should be connected to the A/D converters, via the AA unit.

Coil Out to Sat socket

This connector is the output to the coils in the OSEMs. The cable from this connector is connected to the Satellite box. The signals pass through the Satellite Box then out to the coils via a separate connector on the satellite box.

PD In from Sat socket

This is connected to the satellite box, which generates the amplified Photo Diode (PD) signals. The PD outputs pass straight through the Drive Unit then out to the AA unit.

4.3 Connector Details

PD Out to AA:

Type: 9 Way Socket. Connections: Differential pairs.

Voltage range: +/-10v per line = +/-20v differential.

Calibration: 0.3233V differential per μA photodiode current. 60 μA gives 19.4v differential output.

Connections:

| Pin | Function |
| --- | --- |
| 1 | PD1 Positive |
| 2 | PD2 Positive |
| 3 | PD3 Positive |
| 4 | PD4 Positive |
| 5 | 0v |
| 6 | PD1 Negative |
| 7 | PD2 Negative |
| 8 | PD3 Negative |
| 9 | PD4 Negative |

LED Mon

Type: 9 Way Socket.

Voltage range: 10v.  Nominal output = 1v

Connections:

| Pin | Function |
| --- | --- |
| 1 | Imon 1 Positive |
| 2 | Imon 2 Positive |
| 3 | Imon 3 Positive |
| 4 | Imon 4 Positive |
| 5 | 0v |
| 6 | Imon 1 Negative |
| 7 | Imon 2Negative |
| 8 | Imon 3Negative |
| 9 | Imon 4 Negative |

Drive Inputs

Type: 9 Way Plug. Connections: Differential pairs.

Voltage range: +/-10v per line = +/-20v differential.

Connections:

| Pin | Function |
| --- | --- |
| 1 | Input 1 Positive |
| 2 | Input 2 Positive |
| 3 | Input 3 Positive |
| 4 | Input 4 Positive |
| 5 | 0v |
| 6 | Input 1 Negative |
| 7 | Input 2Negative |
| 8 | Input 3Negative |
| 9 | Input 4 Negative |

Test Inputs

Type: 9 Way Plug, Differential pairs.

Voltage range: +/-10v per line = +/-20v differential.

Note: This connector must be connected to a shorting connector when test are not being carried out.

Connections:

| Pin | Function |
| --- | --- |
| 1 | Test Input 1 Positive |
| 2 | Test Input 2 Positive |
| 3 | Test Input 3 Positive |
| 4 | Test Input 4 Positive |
| 5 | 0v |
| 6 | Test Input 1 Negative |
| 7 | Test Input 2Negative |
| 8 | Test Input 3Negative |
| 9 | Test Input 4 Negative |

Binary IO 1

Type: 37 Way Socket

Single ended with respect to 0v

+12v indicates that the relay is on.

0v indicates that the relay is switched off.

| Pin | Name | Function |
| --- | --- | --- |
| 1 | LP1on1 | Low Pass Filter 1, channel 1 on |
| 2 | 0v |  |
| 3 | 0v |  |
| 4 | Test Mode1 | Test Mode relay 1 on |
| 5 | LP1on2 | Low Pass Filter 1, channel 2 on |
| 6 | 0v |  |
| 7 | 0v |  |
| 8 | Test Mode2 | Test Mode relay 2 on |
| 9 | LP1on3 | Low Pass Filter 1, channel 3 on |
| 10 | 0v |  |
| 11 | 0v |  |
| 12 | Test Mode3 | Test Mode relay 3 on |
| 13 | LP1on4 | Low Pass Filter 1, channel 4 on |
| 14 | 0v |  |
| 15 | 0v |  |
| 16 | Test Mode4 | Test Mode relay 4 on |
| 17 | 0v |  |
| 18 | 0v |  |
| 19 | 0v |  |
| 20 | 0v |  |
| 21 | 0v |  |
| 22 | 0v |  |
| 23 | 0v |  |
| 24 | 0v |  |
| 25 | 0v |  |
| 26 | 0v |  |
| 27 | 0v |  |
| 28 | 0v |  |
| 29 | 0v |  |
| 30 | 0v |  |
| 31 | 0v |  |
| 32 | 0v |  |
| 33 | 0v |  |
| 34 | 0v |  |
| 35 | 0v |  |
| 36 | 0v |  |
| 37 | 0v |  |

Binary IO 2 – Relay Control

Type: 37 Way Plug

Single ended with respect to 0v

0v turns the relay on

Open circuit switches the relay off.

| Pin | Name | Function |
| --- | --- | --- |
| 1 | LP11 | Low Pass Filter 1, channel 1 |
| 2 | 0v |  |
| 3 | 0v |  |
| 4 | Test1 | Test Input 1 |
| 5 | LP12 | Low Pass Filter 1, channel 2 |
| 6 | 0v |  |
| 7 | 0v |  |
| 8 | Test2 | Test Input 2 |
| 9 | LP13 | Low Pass Filter 1, channel 3 |
| 10 | 0v |  |
| 11 | 0v |  |
| 12 | Test3 | Test Input 3 |
| 13 | LP14 | Low Pass Filter 1, channel 4 |
| 14 | 0v |  |
| 15 | 0v |  |
| 16 | Test4 | Test Input 4 |
| 17 | 0v |  |
| 18 | 0v |  |
| 19 | 0v |  |
| 20 | 0v |  |
| 21 | 0v |  |
| 22 | 0v |  |
| 23 | 0v |  |
| 24 | 0v |  |
| 25 | 0v |  |
| 26 | 0v |  |
| 27 | 0v |  |
| 28 | 0v |  |
| 29 | 0v |  |
| 30 | 0v |  |
| 31 | 0v |  |
| 32 | 0v |  |
| 33 | 0v |  |
| 34 | 0v |  |
| 35 | 0v |  |
| 36 | 0v |  |
| 37 | 0v |  |

DC In: Power cable:

Type: 3 pin heavy duty Voltages: Nominally +/- 17volts from Power supply

| Pin | Position | Function |
| --- | --- | --- |
| A1 | Left viewed from rear | Positive |
| A2 | Centre | RTN |
| A3 | Right viewed from rear | Negative |

Noise Monitor

Type: 9 Way Socket. Differential pairs, ac coupled filtered, amplified noise.

| Pin | Function |
| --- | --- |
| 1 | Noise Monitor 1 |
| 2 | Noise Monitor 2 |
| 3 | Noise Monitor 3 |
| 4 | Noise Monitor 4 |
| 5 | 0v |
| 6 | 0v |
| 7 | 0v |
| 8 | 0v |
| 9 | 0v |

V, I and RMS monitor (Pre Production model only)

Type: 25 way socket.  Voltage, Current, and RMS Monitors.

Voltage range: +/-10v per line = +/-20v differential.

| Pin | Name | Function |
| --- | --- | --- |
| 1 | VM4 | Channel 4 Voltage monitor |
| 2 | FC4 | Channel 4 Current monitor |
| 3 | SC4 | Channel 4 RMS monitor |
| 4 | VM3 | Channel 3 Voltage monitor |
| 5 | FC3 | Channel 3 Current monitor |
| 6 | SC3 | Channel 3 RMS monitor |
| 7 | VM2 | Channel 2 Voltage monitor |
| 8 | FC2 | Channel 2 Current monitor |
| 9 | SC2 | Channel 2 RMS monitor |
| 10 | VM1 | Channel 1 Voltage monitor |
| 11 | FC1 | Channel 1 Current monitor |
| 12 | SC1 | Channel1 RMS monitor |
| 13 | 0v |  |
| 14 | 0v |  |
| 15 | 0v |  |
| 16 | 0v |  |
| 17 | 0v |  |
| 18 | 0v |  |
| 19 | 0v |  |
| 20 | 0v |  |
| 21 | 0v |  |
| 22 | 0v |  |
| 23 | 0v |  |
| 24 | 0v |  |
| 25 | 0v |  |

V, I and RMS monitor - Production model

Type: 25 way plug.  Voltage, Current, and RMS Monitors.

Voltage range: +/-10v per line = +/-20v differential.

| Pin | Name | Function |
| --- | --- | --- |
| 1 | SC1 | Channel 1 RMS monitor |
| 2 | FC1 | Channel 1 Current monitor |
| 3 | VM1 | Channel 1 Voltage monitor |
| 4 | SC2 | Channel 2 RMS monitor |
| 5 | FC2 | Channel 2 Current monitor |
| 6 | VM2 | Channel 2 Voltage monitor |
| 7 | SC3 | Channel 3 RMS monitor |
| 8 | FC3 | Channel 3 Current monitor |
| 9 | VM3 | Channel 3 Voltage monitor |
| 10 | SC4 | Channel 4 RMS monitor |
| 11 | FC4 | Channel 4 Current monitor |
| 12 | VM4 | Channel 4 Voltage monitor |
| 13 | 0v |  |
| 14 | 0v |  |
| 15 | 0v |  |
| 16 | 0v |  |
| 17 | 0v |  |
| 18 | 0v |  |
| 19 | 0v |  |
| 20 | 0v |  |
| 21 | 0v |  |
| 22 | 0v |  |
| 23 | 0v |  |
| 24 | 0v |  |
| 25 | 0v |  |

Coil Out to Sat:       Type: 15 way socket.  Drive to the OSEM coils, delivered via the satellite box. Current output, 200mA maximum, differential drive.

| Pin | Name | Function |
| --- | --- | --- |
| 1 | CoilST1 | Coil 1 Start |
| 2 | 0v |  |
| 3 | CoilST2 | Coil 2 Start |
| 4 | 0v |  |
| 5 | CoilST3 | Coil 3 Start |
| 6 | 0v |  |
| 7 | CoilST4 | Coil 4 Start |
| 8 | 0v |  |
| 9 | CoilFN1 | Coil 1 Finish |
| 10 | 0v |  |
| 11 | CoilFN2 | Coil 2 Finish |
| 12 | 0v |  |
| 13 | CoilFN3 | Coil 3 Finish |
| 14 | 0v |  |
| 15 | CoilFN4 | Coil 4 Finish |

PD in from SAT

Type: 25 way socket.  Differential Photodiode outputs from the satellite box

Voltage range: +/-10v per line = +/-20v differential.

Current monitor outputs from the satellite box: 1v nominal

Power Supplies to the Satellite Box: +/- 17v nominal

Connections:

| Pin | Name | Function |
| --- | --- | --- |
| 1 | PD1P | Photodiode output channel 1 Positive |
| 2 | PD2P | Photodiode output channel 2 Positive |
| 3 | PD3P | Photodiode output channel 3 Positive |
| 4 | PD4P | Photodiode output channel 4 Positive |
| 5 | Imon1P | LED Current Monitor 1 Positive |
| 6 | Imon2P | LED Current Monitor 2 Positive |
| 7 | Imon3P | LED Current Monitor 3 Positive |
| 8 | Imon4P | LED Current Monitor 4 Positive |
| 9 | V+ | Satellite Box Raw Positive Supply |
| 10 | V+ | Satellite Box Raw Positive Supply |
| 11 | V- | Satellite Box Raw Negative Supply |
| 12 | V- | Satellite Box Raw Negative Supply |
| 13 | 0v |  |
| 14 | PD1N | Photodiode output channel 1 Negative |
| 15 | PD2N | Photodiode output channel 2 Negative |
| 16 | PD3N | Photodiode output channel 3 Negative |
| 17 | PD4N | Photodiode output channel 4 Negative |
| 18 | Imon1N | LED Current Monitor 1 Negative |
| 19 | Imon2N | LED Current Monitor 2 Negative |
| 20 | Imon3N | LED Current Monitor 3 Negative |
| 21 | Imon4N | LED Current Monitor 4 Negative |
| 22 | 0v | Satellite Box 0v |
| 23 | 0v | Satellite Box 0v |
| 24 | 0v | Satellite Box 0v |
| 25 | 0v | Satellite Box 0v |

5 FUNCTIONAL DESCRIPTION

5.1 Inside the Unit

FIG. 3 Transmon Driver Unit Block Diagram

The Transmon Coil Drive Unit contains a Coil Drive board and a Monitor board. The purpose of the Monitor Board is to monitor the Output voltage, Output Current, RMS Current and Output Noise from unit.

The Transmon Driver Unit also passes the amplified signals from the Photodiodes, which detect the position of the OMC mirror, back to the control electronics without processing them in any way.

5.2 The Drive Board

FIG.4 Transmon Driver Channel Block Diagram

Each Transmon Driver board consists of four identical channels, one of which is shown above. Three power regulators (not shown), which provide regulated power to the four channels, are also mounted on the board.

Taking the diagram block by block, the first block contains relays which switch the circuit between the demand inputs and the test inputs.

The second block contains a low pass filter with a corner frequency of 1 Hz, followed by a complimentary zero at 10 Hz. To a good approximation, the gain is reduced by a factor of 0.7 at 1 Hz, the attenuation increases at a rate of 20dB/decade up to the corner frequency of the zero at 10 Hz, after which the characteristic levels off. This stage has a gain of 1.2. The third block is similar, except that it has a gain of 1, and the filter is not switchable. The overall gain of the TOP driver is 1.2.

The forth block is a driver stage, providing the required 200mA maximum output.

The voltage which will appear across the coil, and the output voltage, are both buffered by unity gain voltage followers before being fed to the monitor circuit.

The first filter may be switched in and out as required by its own control relay, and a talkback signals confirm the positions of relay.

The Test Mode relay provided on the front end enables the amplifier inputs to be switched to the Test input when required.

5.3 Monitor Board

Fig 6 Block diagram of one channel of the monitor board.

Description

The function of the Monitor Board is to monitor the outputs from a drive board. It has four identical channels, one per drive board channel.

Inputs

The inputs of the monitor Board are connected to the two amplifier outputs and the coil feeds.

The following signals are monitored on each channel. They are: the positive and negative output voltages from the drive amplifiers, and the positive and negative output voltages developed across the coil.

These four signals are used to derive the amplifier output voltage, the current through the coil and the output noise on each of the drive amplifier channels.

Output Voltage

The Amplifier Output Voltage is measured by subtracting the Positive Amplifier output from the Negative Amplifier Output. The output is scaled by a factor of 3, so, for example, inputs of +15v and -15v, the sum of which is 30v, would give an output of 30v/3 = 10v.

Coil Current

The coil current is monitored by measuring the sum of the voltages across the two output resistors (R1 and R2 on the block diagram.) The amplifier performs the following calculation:

{(Pos Voltage Output) – (Negative Voltage Output)} -

{(Pos Coil Voltage) - (Negative Coil Voltage)}

The voltage across the coil is subtracted from the voltage across the Amplifier output. This gives the voltage across the output resistors, which is proportional the coil current.

An R.M.S. converter chip calculates the true R.M.S. output current. The overall scaling factor of the R.M.S. converter and amplifier circuits is 1/3.

Noise Measurement

As the noise level amounts to a few Pico amperes, it is extremely difficult to measure directly. Instead, the noise voltage across the amplifier outputs is monitored. This enables the coil noise current to be estimated.  Four amplifier stages each coupled with a high pass filter are used, followed by a two stage low pass filter.

6. Assembly and Disassembly

The unit may be disassembled in the usual way. The lid is removed by removing the screws, revealing the two PCBs and the wiring looms. In order to remove the PCBs, the screw locks must first be removed. To remove the monitor board, the screw holding the support pillar, and the pillars must be removed.

Note that the drive board is held in place by the heat sink screws. These must be removed to extract the board.

To re-assemble thee unit, note that the heat sink bar is isolated from the PCB by a green electrical insulating layer. The easiest method of assembly is to feed the end M3 screws through, place the unit on a table, then place the heat sink over the screws, with the counter sinking on top. Take casreto get the heat sink the right way up, by reference to the power driver IC orientation on the board.

Having got the heat sink in place, gently ease the green insulating sheet in place over the heat sink, ensuring that the heat sink is covered. Then lower the PCB in place. Replace the black insulating boshes, cylindrical side down, then screw on the M3 nuts to hold everything in place. Finally insert the other screws, and add the bushes and nuts.

After assembly it is wise to check that the tabs of the power driver devices are electrically insulated from the case.

6. Assembly and Disassembly

The unit may be disassembled in the usual way. The lid is removed after removing the screws, revealing the two PCBs and the wiring looms. In order to remove the PCBs, the screw locks must first be removed. To remove the monitor board, the screw holding the support pillar, and the pillars must be removed.

Note that the drive board is held in place by the heat sink screws. These must be removed to extract the board.

To re-assemble thee unit, note that the heat sink bar is isolated from the PCB by a green electrical insulating layer. The easiest method of assembly is to feed the end M3 screws through, place the unit on a table, then place the heat sink over the screws, with the counter sinking on top. Take care to get the heat sink the right way up, by reference to the power driver IC orientation on the board.

Having got the heat sink in place, gently ease the green insulating sheet in place over the heat sink, ensuring that the heat sink is covered. Then lower the PCB in place. Replace the black insulating boshes, cylindrical side down, then screw on the M3 nuts to hold everything in place. Finally insert the other screws, and add the bushes and nuts.

After assembly it is wise to check that the tabs of the power driver devices are electrically insulated from the case.

# Text boxes

Test Mode

Relays

LP Filter 1

Fc =1Hz

10 Hz zero

Mode

Control

Filter 1

In/out

LP Filter 2

Fc =1Hz

10 Hz zero

Power

Output

Stage

Monitor

Buffers

O/Ps to Coils

Outputs to Monitor Board

Coil Demand Signals from AI

Coil Drive Electronics

Monitor

Board

O/Ps to Coils

Coil Demand Signals from AI

Vmon

Imon

I r.m.s

Noise

Imonn

R1

Vmon

Vp-Vn

Cp

Vp

R2

Coil

Vn

Cn

(Vp-Cp)+(Cn-Vn)

Noise Monitor

Coil Driver Channel

Digital Control System

D/A

Convertors

AI

Unit

Coil Drive Unit

Coil Drivers

Monitors

OSEM

Coils

Photodiodes

Satellite Box

Photodiode

Amplifiers

Flag                       Magnet

Transmon Stage

AA Unit

D/A

Convertors

Demand

Feedback

PSU

CHANNEL 2

CHANNEL 1

CHANNEL 3

CHANNEL 4

Monitor Board

Drive Board

Power Switch

PD out to AA

LED Mon

Drive Input

Test In

Binary I/O 1

Binary I/O 2

PD in from Sat

Coil out to Sat

Power

Noise Monitor

V, I & RMS Monitor

PSU