## GS13
Issue Tracking for Seismometers and Pod Assemblies
| D1000166 | D047810/D0900857 |  | Issue |  |  |  |  | Action
Type | SN | Pod SN # | H/V | Location | Sub-location | Date Issue Found | Description | Elog/Alog/DCC # | Action Planned | Location | Action Date | Action taken, Diagnosis, Results | Person | Status Pending/Cosed
HAM GS13 | 807? | 46 | H | LHO | HAM-ISI #1, H3 | 2010-06-24 | Power spectrum looked like the instrument is locked below 10 Hz. | LHO aLOG #31, DCC T1000487 | Replace seismometer | LLO | 2010-06-28 | Seismometer was replaced by S/N 68 |  | Closed
[9 empty] | Investigate/solve issue on instrument | LLO | 2010-08-17 | One hanging fixture was found to had come loose, it was retorqued- cf T1000487 |  | Closed
[9 empty] | Huddle Test, Re-pod and Leak test. | LLO | 2012-04-06 | Reseated all BeCu flexures, aligned cocked flex strip cap squares, adjusted the natural frequency, PM test executed. |  | Closed
BSC GS13 | Not recorded | Not recorded. | H | MIT | Dirty Assy, Hor. | 2011-01-27 | Leveling issue. | SEI elog 1881, DCC E1101111 | Open the pod and inspect the flexures | MIT | 2011-01-27 | Was due to a tiny gap between the flexure tip and the clamp. Fixed. See DCC E1101111. |  | Closed
BSC GS13 | 719 | 98 | V | LHO | BSC1, V3 | 2011-06-17 | Grounded Mass, Opened and re-centered. | LHO aLOG 939 | Huddle Test, Re-pod and Leak test. (installed by mistake in Pod 22) | LLO |  |  |  | Pending
BSC GS13 |  | 88 |  | LHO | Removed from BSC2, V1 | 2011-04-12 | Incorrect Length Bolts used for pod flange. | LLO aLOG 1324 | Re-pod and test. | LLO |  | Repod and approved | Ray M | Closed
BSC GS13 |  | 37 |  | LHO | Removed from BSC2, H1 | 2011-04-12 | Incorrect Length Bolts used for pod flange. | LLO aLOG 1324 | Re-pod and test. | LLO |  | Repod and approved | Ray M | Closed
BSC GS13 |  | 52 |  | LHO | Removed from BSC2, V2 | 2011-04-12 | Incorrect Length Bolts used for pod flange. | LLO aLOG 1324 | Re-pod and test. | LLO |  | Repod and approved | Ray M | Closed
BSC GS13 |  | 14 |  | LHO | Removed from BSC2, H2 | 2011-04-12 | Incorrect Length Bolts used for pod flange. | LLO aLOG 1324 | Re-pod and test. | LLO |  | Repod and approved | Ray M | Closed
BSC GS13 |  | 48 |  | LHO | Removed from BSC2, V3 | 2011-04-12 | Incorrect Length Bolts used for pod flange. | LLO aLOG 1324 | Re-pod and test. | LLO |  | Repod and approved | Ray M | Closed
BSC GS13 |  | 27 |  | LHO | Removed from BSC2, H3 | 2011-04-12 | Incorrect Length Bolts used for pod flange. | LLO aLOG 1324 | Re-pod and test. | LLO |  | Repod and approved | Ray M | Closed
HAM GS13 | 723? | 52 | V | MIT | Vert | 2011-09-26 | Weird low freq. behavior |  | Examine behavior and inspect mechanical. (installed by mistake in Pod 60 in BSC Unit 2) | LLO |  |  |  | Pending
HAM GS13 |  | 68 |  | LHO | Removed from BSC 1 | 2011-09-30 | Incorrect Length Bolts used for pod flange. |  | Re-pod and test. | LLO |  | Repod and approved | Ray M | Closed
HAM GS13 |  | 76 |  | LHO | Removed from BSC 1 | 2011-09-30 | Incorrect Length Bolts used for pod flange. |  | Re-pod and test. | LLO |  | Repod and approved | Ray M | Closed
HAM GS13 |  | 84 |  | LHO | Removed from BSC 1 | 2011-09-30 | Incorrect Length Bolts used for pod flange. |  | Re-pod and test. | LLO |  | Repod and approved | Ray M | Closed
HAM GS13 |  | 85 |  | LHO | Removed from BSC 1 | 2011-09-30 | Incorrect Length Bolts used for pod flange. |  | Re-pod and test. | LLO |  | Repod and approved | Ray M | Closed
HAM GS13 |  | 89 |  | LHO | Removed from BSC 1 | 2011-09-30 | Incorrect Length Bolts used for pod flange. |  | Re-pod and test. | LLO |  | Repod and approved | Ray M | Closed
BSC GS13 | 740 | 77 | V | LHO | ? | 2011-09-30 | Dropped from about 1 foot. | LHO alog 1473. | Must be inspected. | LLO | 2011-10-26 | Broken flexures+ broken tag+ loose nuts |  | Closed
[8 empty] | LLO alog 1958 | Replace broken flexures | LLO | 2011-10-27 | Done
[8 empty] | LLO alog 1958 | Change washers | LLO | 2011-11-16 | Stilll haven't found the right washer thickness
[8 empty] | LLO alog 1958 | Replace tag-Send to Geotech? | LLO | 2012-04-06 | Replaced broken proof mass to cantilever flex strip with replacement from Stanford, reseated all BeCu flexures, fixed oil canned flex strips, adjusted the natural frequency, PM test executed. | Dan C
Huddle Test, Re-pod and Leak test.
BSC GS13 | 781 | 103 | H | LHO | ? | 2011-09-30 | Dropped from about 1 foot. | LHO alog 1473. | Must be inspected. | LLO | 2011-11-16 | No visible damage |  | Closed
[8 empty] | LLO alog 1958 | Replace broken flexures | LLO | 2011-10-27 | Done
[8 empty] | LLO alog 1958 | Change washers | LLO | 2011-11-16 | Stilll haven't found the right washer thickness
[9 empty] | Huddle Test, Re-pod and Leak test. | LLO | 2012-04-06 | Proof mass was off center and BeCu flexures were stressed. Recentered the mass by reseating all of the BeCu flexures. Adjusted the natural frequency, PM test executed. | Dan C
GS13 | 696 | N/A | V | LLO | Geotech |  | Failed ARHT-No response/coherence with STS 2 | SeismicSVN/Common/Data/aLIGO_GS13_TestData/AsReceived_TestResults_PDFs/LV696_AR_HuddleTF.pdf | Vendor rework | Geotech | 2011-11-16 | At vendor |  | Pending
GS13 | 863 | N/A | H | LLO | Geotech | 2011-09-28 | Failed ARHT-No response/coherence with STS 2 | SeismicSVN/Common/Data/aLIGO_GS13_TestData/AsReceived_TestResults_PDFs/LH863_AR_HuddleTF.pdf | Vendor rework | Geotech | 2011-11-16 | At vendor |  | Pending
GS13 | 697 | N/A | V | LLO | SPSR | 2010-12-02 | Failed PMQA-can not get rid of 105 Hz resonance | SeismicSVN/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_PDFs/LV697_PM_DrivenTF.pdf | Troubleshoot o-ring | LLO | 2012-04-06 | Released tension in all BeCu flexures and torsion rod, reseated all BeCu flexures, fixed “oil canning” on several flex strips, adjusted the natural frequency, PM test executed. | Dan C | Closed
GS13 | 708 | N/A | V | LLO | SPSR | 2010-12-02 | Failed PMQA- can not lower the resonant frequency to 1 Hz | SeismicSVN/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_PDFs/LV708_PM_HuddleTF.pdf | Troubleshoot adjustment rod | LLO |  |  |  | Pending
GS13 | 716 | N/A | V | LLO | SPSR | 2011-11-11 | Failed PMQA-can not get rid of 105 Hz resonance | SeismicSVN/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_PDFs/LV716_LV716_PM_HuddleTF.pdf | Troubleshoot o-ring | LLO | 2012-04-06 | Released tension in all BeCu flexures and torsion rod, realigned flex strip from the end of cantilever bar to hang vertically, reseated all BeCu flexures, adjusted the natural frequency, PM test executed. | Dan C | Closed
GS13 | 720 | N/A | V | LLO | SPSR | 2011-08-25 | Failed PMQA- can not increase the resonant frequency to 1 Hz | SeismicSVN/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_PDFs/LV720_PM_HuddleTF.pdf | Troubleshoot adjustment rod | LLO | 2012-04-06 | Released tension in all BeCu flexures and torsion rod, reseated all BeCu flexures, adjusted the natural frequency, PM test executed. | Dan C | Closed
GS13 | 786 | N/A | H | LLO | SPSR | 2010-12-02 | Failed - no response/coherence with STS 2 | SeismicSVN/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_PDFs/LH786_PM_HuddleTF.pdf | Troubleshoot Mods. | LLO | 2012-04-06 | The pre-amp board S1000090 had a discolored negative supply filter capacitor. After hooking up the power supply it was noted that it was drawing 4 amps. This is due to an incorrectly GeoTech wired cable which had swapped vcc- and ground. Swapping the #13 and #15 pins on J1 solved the problem and a new pre-amp board was installed (new board s/n S1105438). PM test executed. | Dan C | Closed
GS14 | 837 | N/A | H | LLO | SPSR | 2011-02-24 | Failed PMQA-noisy! | SeismicSVN/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_PDFs/LH837_PM_HuddleTF.pdf | Installed by mistake in pod 61 (BSC Unit 2) | LLO |  |  |  | Pending
GS13 | 805 | 68 | H | LHO | Staging Building | 2011-11-28 | The Mass seems to be stuck | LHO alog 1801 and 1822 | POD open at LHO and reuse for testing | LHO | 2012-04-06 | All three upper BeCu flexure to proof mass attachment fittings were loose. One lower flexure was broken (perhaps as a result of the loose upper attachments?) All attachments tightened, lower BeCu flexure replaced, all flexures reseated, torsion rod recentered, adjusted the natural frequency, PM test executed. | Dan C | Closed
|  |  |  | LLO | SPSR | 2012-01-06 | The Mass is not moving freely | LHO alog 1997 | POD open at LHO and will not be reused for testing. Will be send back to LLO | LLO
HAM GS13 | 798 | 38 | H | LLO | LVEA | 2012-01-26 | Lower response below 15 Hz | LLO alog 2353 | Change flexure? | LLO | 2012-04-06 | One loose upper BeCu flexure to proof mass attachment was found. Tightened all three upper attachments, reseated the BeCu flexures, adjusted the natural frequency, PM test executed. | Dan C | Closed
HAM GS13 | 800 | 94 | H | LHO | Staging Building | 2012-03-01 | Gain of approx. 1/2 | LHO alog 2321 and 2365. Inspection/reparation in LHO aLog 2369 | Intrument will be made vacuum compatible at LLO | LHO | 2012-03-06 | Preamplifier and socket-board (taken from pod #68) replaced. Gain OK | Greg G | Closed
[10 empty] | LLO | 2012-04-07 | PM test executed | Ian M
HAM GS13 | 779 | 71 |  | LHO | Staging Building | 2012-03-01 | Power Spectrum looks like the instrument is locked below 20Hz | LHO alog 2321 and 2365. Inspection/reparation in LHO aLog 2369 | Intrument will be opend, inspected at LHO. Instrument will be repaired at LLO | LHO | 2012-03-06 | One broken flexure. Flexure replaced. The instrument readout is now a constant -32000counts | Greg G | Pending
[10 empty] | LHO | 2012-03-14 | Shipping Back to LLO | Mitchell R
HAM GS13 | 794 | 24 | H | LHO | Staging Building | 2012-03-15 | Power Spectrum looks like the instrument is locked below 20Hz | LHO aLog #2449 and # 2471 | This intrument will be opend, inspected and repaired at LLO | LHO | 2012-03-22 | The pod was opened at LHO. Nothing abnormal was found. Leveling the instrument on the test bench did not affect its response | Greg G | Closed
[10 empty] | LLO | 2012-04-06 | All three upper BeCu flexure to proof mass attachment fittings were loose. The flexures themselves did not have to be released to discover this. Simply releasing the lock slightly allowed the proof mass to be moved by hand. All attachments tightened, reseated BeCu flexures and torsion rod, adjusted the natural frequency, PM test executed | Dan C | Closed
HAM GS13 | 802 | N/A | H | LLO | SPSR | 2010-11-01 | Sensor foot galled in leveling foot | N/A | Change foot | LLO | 2012-04-06 | Replaced galled foot attachment rod with spare replacement from Stanford, installed pre-amp S1105448, reseated all BeCu flexures, adjusted the natural frequency, PM test executed. | Dan C | Closed
GS-13 | 784 | N/A | H | LLO | SPSR |  | No know issues |  |  | LLO | 2012-04-06 | Everything internal looked normal - PM test executed. | Dan C | Closed
GS-13 | 785 | N/A | H | LLO | SPSR |  | No know issues |  |  | LLO | 2012-04-06 | Centered torsion rod, released stress in rod, reseated BeCu flexures, adjusted the natural frequency, PM test executed. | Dan C | Closed
GS-13 | 796 | N/A | H | LLO | SPSR |  | Need Post Modified testing |  |  | LLO | 2012-04-06 | Centered torsion rod, released stress in rod, reseated BeCu flexures, PM test executed. | Dan C | Closed
GS-13 | 815 | N/A | H | LLO | SPSR |  | No record |  |  | LLO | 2012-04-06 | In this seismometer the BeCu flexures are visibly not straight. This mismatch in the pads that the flexure spans from the structure to the proof mass is likely caused by a slight machining error in manufacturing. This does not seem to affect performance. Centered torsion rod, released stress in rod, reseated BeCu flexures, adjusted the natural frequency, PM test executed. | Dan C | Closed
GS-13 | 712 | N/A | V | LLO | SPSR |  | Can and instrument itself do not seem to fit properly -> "strange lip" |  |  | LLO | 2012-04-06 | Reported “strange lip” investigated. This is caused by a mis-machined can and should not affect the nstrument’s performance. Checked internal flexures and flex strips, PM test executed. | Dan C | Closed
GS-13 | 741 | N/A | V | LLO | SPSR |  | Could not adjust o-ring to damp spring resonance |  |  | LLO | 2012-04-06 | Adjusted the mass position. PM test executed | Dan C | Closed
GS-13 | 747 | N/A | V | LLO | SPSR |  | Unable to adjust resonant frequency |  |  | LLO | 2012-04-06 | Released tension in all BeCu flexures and torsion rod, reseated all BeCu flexures, adjusted the natural frequency, PM test executed | Dan C | Closed
GS-13 | 761 | N/A | V | LLO | SPSR |  | Bent torsion rod (resonant frequency adjustment) |  |  | LLO | 2012-04-06 | Released tension on torsion rod, realigned flex strip cap squares, reseated BeCu flexure, adjusted the natural frequency, PM test executed. | Dan C | Closed
GS-13 | 766 | N/A | V | LLO | SPSR |  | No record |  |  | LLO | 2012-04-06 | Readjusted all cocked flex strip hold down hardware, reseated all BeCu flexures with torsion rod centered, adjusted the natural frequency, PM test executed | Dan C | Closed
GS-13 | 812 | 18 | H | LHO | Staging Building | 2012-04-11 | Power Spectrum looks like the instrument is locked below 20Hz | LHO aLog #2598 | Shipping back to LLO for opening and inspection | LHO | 2012-04-19 | Sent Back to LLO | Mitchell R | Pending
GS-13 | 827 | 20 | H | LHO | Staging Building | 2012-04-11 | Does not work consistantly. Power Spectrum looks like the instrument is locked below 20Hz when doesn not work | LHO aLog #2598 | Installed on HAM-ISI #5 for testing of the unit and sent back to LLO | LHO | 2012-05-21 | Sent Back to LLO | Mitchell R | Pending
GS-13 | 845 | 27 | H | LHO | Staging Building | 2012-04-11 | Does not work consistantly. Power Spectrum looks like the instrument is locked below 20Hz when doesn not work | LHO aLog #2598 | Installed on HAM-ISI #5 for testing of the unit and sent back to LLO | LHO | 2012-05-21 | Sent Back to LLO | Mitchell R | Pending
GS-13 | 830 | 37 | H | LHO | Staging Building | 2012-04-11 | Does not work consistantly. Power Spectrum looks like the instrument is locked below 20Hz when doesn not work | LHO aLog #2598 | Installed on HAM-ISI #5 for testing of the unit and sent back to LLO | LHO | 2012-05-21 | Sent Back to LLO | Mitchell R | Pending
GS-13 | 847 | 84 | H | LHO | Staging Building | 2012-04-11 | Does not work consistantly. Power Spectrum looks like the instrument is locked below 20Hz when doesn not work | LHO aLog #2598 | Installed on HAM-ISI #5 for testing of the unit and sent back to LLO | LHO | 2012-05-21 | Sent Back to LLO | Mitchell R | Pending
GS13 | 833 | 94 | H | LHO | Staging Building | 2012-06-20 | Did not work after being installed in the ISI. Power Spectrum looks like the seismic mass is stuck below 60Hz | LHO aLogs 3183 | Open Pod For Inspection | LHO | 404861 | Pod opened. Two broken flexures and debris were found. See LHO aLog #3198 | HugoP GregG | Pending
[9 empty] | Replace broken flexures | LHO | 2012-07-02 | Replaced 2 flexures. Re-podded the sensor. Spectra of the instrument in the pod showed instrument failure: we suppose it has its shaft bent.
[9 empty] | Shipped to LLO for replacement of the seismometer | LHO | 2012-08-07 | up to LLO | Mitchell R

## L4C
Issue Tracking for Seismometers and Pod Assemblies
|  |  | D047820 | Issue |  |  |  |  | Action
Type | S/N | H/V | Pod SN # | Location | Sub-location | Date | Description | Elog/Alog/DCC # | Action Planned | Location | Action Date | Action taken, Diagnosis, Results | Person | Status Pending/Closed
L4C | H 1097 | H | 34 | LHO | Removed from BSC2, H1 |  | Not sure of neon leak testing record. |  | Re-neon leak testing | LLO | 2011-10-17 | Passed neon leak testing |  | Closed
L4C | 937 | V | x-34 | MIT | Removed from HAM |  | No neon leak testing record. |  | Re-neon leak test. | LLO | TBD |  |  | Closed
L4C | 1072 | V | x-45 | LHO | Removed from BSC2. |  | Non Helium Leak Tested chamber. |  | Do not use in aLIGO vacuum |  |  |  |  | Closed
L4C | H 959 | H | 10 | LHO | Removed from BSC2. |  | Pressure Sensor level incorrect. |  | Replace Socket Board. | LLO | 2011-09-19 | Replaced S1103521 with S1103925. |  | Closed
L4C | 963 | H | X-24 | MIT |  | 2011-09-22 | Weak Signal, pre-amp failed? |  | Replace Socket Board. | LLO |  |  |  | Pending
L4C | 1090 | V | X-26 | MIT |  | 2011-09-22 | Weak Signal, pre-amp failed? |  | Replace Socket Board. | LLO |  |  |  | Pending
L4C | H 969 | H | N/A | LLO | SPSR | 2011-11-16 | Failed ARHT-2 Hz resonance |  | Vendor Rework | Sercel | 2011-11-23 |  |  | Closed
[10 empty] | LLO | 2012-03-16 | Passed receiving test | Ian M
L4C | H 1100 | H | N/A | LLO | SPSR | 2011-11-16 | Failed ARHT-2 Hz resonance |  | Vendor Rework | Sercel | 2011-11-23 |  |  | Closed
[10 empty] | LLO | 2012-03-16 | Passed receiving test | Ian M
L4C | H 1117 | H | N/A | LLO | SPSR | 2011-11-16 | Failed ARHT-2 Hz resonance |  | Vendor Rework | Sercel | 2011-11-23 |  |  | Closed
[10 empty] | LLO | 2012-03-16 | Passed receiving test | Ian M
L4C | L41375 |  | N.A | LHO | Removed from HEPI | 2011-10-01 | strange behavior at 1Hz | [6 empty] | Pending
L4C | 919 | V | N/A | LLO | SPSR | 2011-11-11 | Failed ARHT-high resonance | [6 empty] | Pending

## Trilliums
Issue Tracking for Seismometers and Pod Assemblies
|  | D0900648 | Issue |  |  |  |  | Action
Type | S/N | Pod SN # | Location | Sub-location | Date Issue Found | Description | Issue decription: Alog, Elog/DCC # | Action Planned | Location | Action Date | Action taken, Diagnosis, Results | Status Pending/Cosed
T240 OBS | 104 | 24 | LHO | BSC | 2011-09-02 | Pressure Sensor Fix needed. |  | Swap pressure sensor board+Neon fill+ leak testing | LLO | 2011-10-25 | Replaced, Re-podded and Passed Ne Test. | Closed
T240 OBS | 109 | 36 | LHO | BSC | 2011-09-02 | Pressure Sensor Fix needed. |  | Swap pressure sensor board+Neon fill+ leak testing | LLO | 2011-10-13 | Replaced, Re-podded and Passed Ne Test. | Closed
T240 OBS | 115 | 1 | LHO | BSC | 2011-09-02 | Pressure Sensor Fix needed. |  | Swap pressure sensor board+Neon fill+ leak testing | LLO | 2011-10-25 | Replaced, Re-podded and Passed Ne Test. | Closed
T240 OBS | 117 | 15 | LHO | BSC | 2011-09-02 | Pressure Sensor Fix needed. |  | Swap pressure sensor board+Neon fill+ leak testing | LLO | 2011-10-25 | Replaced, Re-podded and Passed Ne Test. | Closed
T240 OBS | 133 | 26 | LHO | BSC | 2011-09-02 | Pressure Sensor Fix needed. |  | Swap pressure sensor board+Neon fill+ leak testing | LLO | 2011-10-13 | Replaced, Re-podded and Passed Ne Test. | Closed
T240 OBS | 138 | 12 | LHO | BSC | 2011-09-02 | Pressure Sensor Fix needed. |  | Swap pressure sensor board+Neon fill+ leak testing | LLO | 2011-10-13 | Replaced, Re-podded and Passed Ne Test. | Closed
T240 OBS | 112 | 32 | LHO | BSC 1 | 2011-09-02 | Pressure Sensor Fix needed. | na | Replace with fixed Pressure Sensor. | LLO | 2011-09-19 | First New Pressure sensorfailed, replaced with a new fixed one. Re-podded and passed Ne test. | Closed
T240 OBS | 101 | 19 | LHO | BSC 1 | 2011-09-02 | Pressure Sensor Fix needed. | na | Replace with fixed Pressure Sensor. | LLO | 2011-09-19 | Replaced, Re-podded and Passed Ne Test. | Closed
T240 OBS | 118 | 20 | LHO | BSC 1 | 2011-09-02 | Pressure Sensor Fix needed. | na | Replace with fixed Pressure Sensor. | LLO | 2011-09-19 | Replaced, Re-podded and Passed Ne Test. | Closed
T240 OBS | na | 12 | LHO | na | 2011-09-30 | Dropped. | LHO alog 1473 | Must be inspected. Will be tested in house. Testing planned defined and approved per Nonometric. See DCC # E1101110 | LLO | Pending | Pending | Pending
T240 OBS | 116 | N/A | LLO | na | 2011-09-28 | High frequency (over 4.5 Hz) noise | na | Vendor rework | Nanometrics | 2011-10-04 | Sent back to nanometric. They identified a faulty capacitor plate jumper wire on channel W. Unit is Repaired. See Elog 1972/ DCC E1101112. | Closed
T240 OBS | 140 | N/A | LLO | na | 2011-09-28 | After podding-Can not center anymore(W axis) | na | Pod disassembly, vendor rework | Nanometrics | 2011-10-04 | Sent back to nanometric. They identified a faulty channel W axis as the source of the issue. Unit is Repaired. See Elog 1972/ DCC E1101112. | Closed
T240 OBS | na | 31 | LHO | BSC6 | 2011-10-18 | Centering issue - Intermitent problem | LHO alog 1594 | Will be removed from BSC 6. Will be tested in house. Testing planned defined and approved per Nonometric. See DCC # E1101110 | LLO | Pending | Pending | Pending
T240 OBS | 123 | 3 | LHO | BSC | 2011-10-18 | Flask - POD - holes not aligned |  | Will be sent to LLO | LLO | 2012-03-09 | The problem was found to come from the D0900650. The 1/4" through holes on the outer edge were drilled undersized. Base plate will be modified in-house. | Pending
T240 OBS | 109 | 36 | LHO | BSC | 2011-10-25 | Flask - POD - holes not aligned |  | Will be sent to LLO | LLO | 2012-03-09 | The problem was found to come from the D0900650. The 1/4" through holes on the outer edge were drilled undersized. Base plate will be modified in-house. | Pending
T240 | 528 |  | LASTI | Ground | 2011-10-20 | Intermitent centering process functionning |  |  | LLO
T240 | 516 |  | LASTI | Ground | 2011-10-20 | received from Livingston to replace Trillium #1 |  |  | LLO
T240 OBS | 111 | 1 | LLO | BSC Unit 1 | 2012-02-27 | Could not center the mass+ pressure sensor |  | We confirmed that we can not center the mass of the instrument (even outside the pod) -> send to Nanometrics? | LLO | Pending | Pending | Pending
T240 OBS | 104 | 24 | LLO | BSC Unit 1 | 2012-02-27 | Pressure Sensor Fix needed. (S1103373) |  | Replace with fixed Pressure Sensor. | LLO | 2012-03-01 | Changed board in pod | Closed
T240 OBS | 148 | 17 | LLO | BSC Unit 1 | 2012-03-06 | Pressure Sensor Fix needed. (S1106374) |  | Replace with fixed Pressure Sensor. | LLO | 2012-03-01 | Changed board in pod | Closed
T240 OBS | 138 | 12 | LLO | BSC Unit 1 | 2012-03-06 | Pressure Sensor Fix needed. (S1103956) |  | Replace with fixed Pressure Sensor. | LLO | 2012-03-01 | Changed board in pod | Closed

## STS-2
Issue Tracking for Seismometers and Pod Assemblies
Critical issues are hihlighted in orange.
Type | S/N | Pod SN # | Location | Sub-location | Date Issue Found | Issue | Elog/Alog/DCC # | Action Planned | Action Date | Action taken, Diagnosis, Results | Status Pending/Cosed
STS2 | 100144 | na | LASTI | GROUND | 2011-10-15 | Bad coherence below 0.1 Hz compared with the other STS2 |  | replaced sent to Livingston
STS2 | 89941 | na | LASTI | GROUND | 2011-10-15 | Received from Livingston to replace the bad streckheisn #1
STS2 | 89922 |  | LASTI |  | 2012-04-01 | Bad Low frequency coherence |  | returned to Quanterra
STS2 | 100145 |  | LASTI |  | 2012-07-10 | Bad Low frequency coherence |  | REPAIRED

## CPS
Issue Tracking for CPS
Type | Sensor S/N | Board S/N | Date Issue Found | Issue | SVNElog/Alog/DCC # | Action Planned | Action Date | Action taken, Diagnosis, Results | Status Pending/Cosed
Coarse | 12865 | 12337 | 2011-10-11 | Stopped working - Probably hit during recabling
Coarse | 13062 | 13620 | April 9 -2012 | Initialy Installed on LHO Unit 2 - BSC6 - Removed on April 9, 2012 after cartridge installation - Not working at all - Broken connection at the sensor head - Swapping board doesn't do anything.
Fine | 15865 | 13573 | April 9 -2012 | Initialy Installed on LHO Unit 2 - BSC6 - Removed on April 9, 2012 after cartridge installation - Not working at all - Broken connection at the sensor head - Swapping board doesn't do anything.
Coarse | 15860 | 13577 | April 12 -2012 | Initialy Installed on LHO Unit 2 - BSC6 - Removed on April 12, 2012 after cartridge installation - Noise HF - Grounding issue? | BSC-ISI/H2/ETMY/Data/Misc/CPS_Noise_Investigation_2012_04_11
Fine | 12427 | 12901 | April 12 -2012 | Initialy Installed on LHO Unit 2 - BSC6 - Removed on April 12, 2012 after cartridge installation - Noise HF - Grounding issue? | BSC-ISI/H2/ETMY/Data/Misc/CPS_Noise_Investigation_2012_04_11