*LIGO Laboratory / LIGO Scientific Collaboration*

LIGO- E1101223-v1 *Advanced LIGO* 22 December 2011

Test Procedure for Slow Controls Concentrator Auxiliary 1

Daniel Sigg

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

|  |  |
| --- | --- |
| **California Institute of Technology****LIGO Project – MS 18-34****1200 E. California Blvd.****Pasadena, CA 91125**Phone (626) 395-2129Fax (626) 304-9834E-mail: info@ligo.caltech.edu | **Massachusetts Institute of Technology****LIGO Project – NW22-295****185 Albany St****Cambridge, MA 02139**Phone (617) 253-4824Fax (617) 253-7014E-mail: info@ligo.mit.edu |
| **LIGO Hanford Observatory****P.O. Box 1970****Richland WA 99352**Phone 509-372-8106Fax 509-372-8137 | **LIGO Livingston Observatory****P.O. Box 940****Livingston, LA 70754**Phone 225-686-3100Fax 225-686-7189 |

http://www.ligo.caltech.edu/

# Overview

The slow controls concentrator auxiliary 1 supports 4 analog inputs, 4 analog outputs, 4 binary inputs and 4 binary outputs.

# Test Equipment

* Multimeter
* Second slow controls concentrator auxiliary 1 ([D1102065-v1](https://dcc.ligo.org/DocDB/0074/D1102065/001/D1102065-v1.pdf))
* DC power supplies

# Documentation

* Schematic—[D1102065-v1](https://dcc.ligo.org/DocDB/0074/D1102065/001/D1102065-v1.pdf)

# Tests

Power up the measurement equipment and connect open the lid of the DUT. Connect a DB37 cable (male-male) between the DUT and the second slow controls concentrator. Equip the BNC inputs of the second slow controls concentrator with 50Ω terminators.

## Power

Check the voltages on the concentrator power board. The voltage should be within 5% of nominal.

TP6 (+5V)\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

TP8 (+15V)\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

TP3 (−15V)\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Test that the OK signal is a TTL low (<0.8V).

TP9 (OK) \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

## LED

Check that the LED on the front panel and the 2 LEDs on the rear panel are lit.

Front panel LED\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ Rear panel LEDs\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

## Testing

Use an Ohmmeter and check the continuity of the signal lines between the two slow controls concentrators. Each BNC signal should read 50Ω.

|  |  |  |
| --- | --- | --- |
| **Concentrator** | **Signal** | **Pass/Fail** |
| AI 1 | Auxiliary analog input |  |
| AI 2 | Auxiliary analog input |  |
| AI 3 | Auxiliary analog input |  |
| AI 4 | Auxiliary analog input |  |
| AO 1 | Auxiliary analog output |  |
| AO 2 | Auxiliary analog output |  |
| AO 3 | Auxiliary analog output |  |
| AO 4 | Auxiliary analog output |  |
| BI 1 | Auxiliary binary input |  |
| BI 2 | Auxiliary binary input |  |
| BI 3 | Auxiliary binary input |  |
| BI 4 | Auxiliary binary input |  |
| BO 1 | Auxiliary binary output |  |
| BO 2 | Auxiliary binary output |  |
| BO 3 | Auxiliary binary output |  |
| BO 4 | Auxiliary binary output |  |