LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1200104 	                 LIGO                   March 1st, 2012

aLIGO HAM-ISI, Installation Test Report, Phase II,

LLO HAM 3

E1200104-V6

Joe Hanson, Celine Ramet, Michael Vargas, Adrien Le Roux, Fabrice Matichard

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

Table of contents:

I.	SIDE CHAMBER TESTING	4

	Step 1: GS13	4

	Step 1.1 – Horizontal GS-13s	4

	Step 3: Check level of Stage 1 Optical Table	6

	Step 4: Shim thickness	7

	Step 5: Blade spring profile	7

	Step 6: Gap checks on actuators-after installation on Stage 1	8

	Step 7: Mass budget	8

	Step 8: Lockers adjustment	10

	Step 9 - Electronics Inventory	10

	Step 10 – Cables inventory – E1100822	11

	Step 11 - Set up sensors gap	12

	Step 12 - Measure the Sensor gap	12

	Step 13 - Check Sensor gaps after the platform release	13

	Step 14– Performance of the limiter	13

	Step 14.1 - Test Nº1 - Push “in the general coordinates”	13

	Step 15 - Position Sensors unlocked/locked  Amplitude Spectral Densities	14

	Step 16 - GS13 ASD -tabled tilted	17

	Step 17- GS13 pressure readout	19

	Step 18 - Actuators Sign and range of motion (Local drive)	19

	Step 19 - Static Testing (Tests in the local basis)	20

	Step 20 - Linearity test	21

	Step 21 Cartesian Basis Static Testing	22

	Step 22- Frequency response	23

	Step 22.1 - Local to local measurements	23

	Step 23 - Lower Zero Moment Plane	25

II.	INITIAL IN CHAMBER TESTING	26

	Step 1: Check level of Stage 1 Optical Table	26

	Step 2: Shim thickness	27

	Step 5: Blade spring profile	27

	Step 7: Mass budget	28

	Step 8: Lockers adjustment	30

	Step 10 – Cables inventory – E1100822	30

	Step 11 - Set up sensors gap	30

	Step 13 - Check Sensor gaps after the platform release	31

	Step 15 - Position Sensors unlocked/locked  Amplitude Spectral Densities	31

	Step 16 - GS13 ASD -tabled tilted	34

	Step 17- GS13 pressure readout	35

	Step 18 - Actuators Sign and range of motion (Local drive)	35

	Step 20 - Linearity test	36

	Step 22- Frequency response	37

	Step 22.1 - Local to local measurements	37

	Step 23 - Lower Zero Moment Plane	39

Introduction

This document presents the tests performed to characterize and validate the “HAM-ISI LLO Unit #3”. This unit was the 3rd unit assembled for aLIGO at LLO. This unit was partially assembled in Fall 2010, but following the discovery of unauthorized repairs in the parts, the assembly was interrupted. All parts in questions were disassembled. This unit was the 1st to be re-assembled when the assembly started back in Spring 2011.

This was the first unit pulled out from storage for installation at Livingston in January 2012, to be installed in HAM 3 chamber. Horizontal seismometers, springs tuned mass dampers, and final cables were installed. Also, the Capacitive Position Sensors cables were shielded.

This document is thought as the follow-up of the previous test document, written after assembly. Only modifications and basic functionalities were tested at that time.

There are 4 distinctive parts:

- Part I: side chamber testing results

- Part II: post insertion testing results

- Part III: Transfer functions taken during optics installation 
	3.1 transfer functions after first triple install 
	3.xx ....

- Part IV: Tests in final configuration

# SIDE CHAMBER TESTING

## Step 1: GS13

All the data related to GS-13 post podding testing can be found in the SVN at : SeismicSVN\seismic\Common\Data\aLIGO_GS13_TestData\PostMod_TestResults_PDFs.

E1000058 spreadsheet provides the status of each individual GS-13 at LLO site during aLIGO HAM assembly

Data files in SVN at:

/opt/svncommon/seisvn/seismic/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_RawASCII

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/Common/MatlabTools

- gs13qatest.m

Figures in SVN at:

/opt/svncommon/seisvn/seismic/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_PDFs

## Step 1.1 – Horizontal GS-13s

Huddle testing

Figure  - Huddle testing of Horiz GS-13 813,798, and 780 after aLIGO modifications

## Step 3: Check level of Stage 1 Optical Table

Issues/difficulties/comments regarding this test:

Initially this test was failed, but we adjusted the shim thickness to pass this test (see following step for new locker shims).

Optical Level measurement of Stage 1 at large (12 - 15) number of points.

The table here does not appear to be really tilted. Most of the difference can be included within the .005” flatness of the table.

Acceptance Criteria

- The maximum angle of the table with the horizontal mustn’t exceed  ~100µrad

Test result:	Passed:   X   	Failed:         .

## Step 4: Shim thickness

Issues/difficulties/comments regarding this test: Those shims were adjusted -in comparison to Phase I-to meet level requirements.

| Lockers | Shim thickness (mil) |
| --- | --- |
| A | 130 |
| B | 125 |
| C | 120 |
| D | 125 |

Table – Shims Thickness

Acceptance Criteria

- Inventory is complete

Test result:	Passed:   X   	Failed:         .

## Step 5: Blade spring profile

Because we adjusted the shims, we retook this measurement:

| Blade # | Base (") | Tip(") | Flatness (mils) |
| --- | --- | --- | --- |
| 1 | .4982 | .4909 | +7.3 |
| 2 | .5024 | .4890 | +13.4 |
| 3 | .4922 | .4881 | +4.1 |

Table 1 - Blade profile

Acceptance Criteria:

- Blades must be flat within 0.015" inches.

Test result:	Passed:   X   	Failed:         .

## Step 6: Gap checks on actuators-after installation on Stage 1

The actuator positions were also adjusted.

| Actuator | Front Gap (1/1000") | Back Gap (1/1000") |
| --- | --- | --- |
| V1 | 85 | 80 |
| V2 | 80 | 85 |
| V3 | 90 | 80 |
| H1 | 85 | 85 |
| H2 | 85 | 85 |
| H3 | 85 | 85 |

Acceptance Criteria

- Gaps must be within 0.010” of design (i.e. 0.090” and .070” pass, but 0.095” and 0.065” doesn’t).

Test result:	Passed:   X   	Failed:         .

## Step 7: Mass budget

Figure – Keel Masses and Wall masses location

Figure  - Optical table masses distribution

|  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| w9 | 1 | 1 |  |  |  | 1 |  | 17.3 | 7.85 |
| w1 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w2 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w3 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w4 |  | 2 |  | 1 |  | 1 | 1 | 49.5 | 22.45 |
| w5 | 1 | 2 |  |  |  | 1 | 1 | 45.6 | 20.68 |
| w6 |  | 3 |  |  |  | 1 | 1 | 46.1 | 20.91 |
| w7 |  | 3 |  |  |  | 1 | 1 | 46.1 | 20.91 |
| w8 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| Side Masses Total | 2 | 11 | 0 | 1 | 0 | 9 | 8 | 375.8 | 170.46 |

Table – Wall masses distribution

|  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| k1 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k2 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
| k3 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k4 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
| k5 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k6 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
|  | 3 | 3 | 3 | 0 | 3 | 0 | 6 | 198.6 | 90.08 |

Table – Keel masses distribution

|  | D972213 | D972215 | D0901075 |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | 2.5 kg | 5 kg | 10 kg |  |  |
|  | 610 | 230 | 5.5 | 11 | 22 | lbs | kgs |
| A | 1 |  |  | 1 |  | 621 | 281.68 |
| B |  |  |  | 1 |  | 11 | 4.99 |
| C |  |  |  | 1 |  | 11 | 4.99 |
| D |  |  |  | 1 |  | 11 | 4.99 |
| E-1 |  |  |  | 1 |  | 11 | 4.99 |
| E-2 |  |  |  | 1 |  | 11 | 4.99 |
| E-3 |  |  |  | 1 |  | 11 | 4.99 |
| Top Masses | 1 | 0 | 0 | 7 | 0 | 687 | 311.62 |

Table - Optical Table Masses distribution

|  | Side | Keel | Top | Total |
| --- | --- | --- | --- | --- |
| Weigh (kg) | 170.46 | 90.08 | 311.62 | 572.16 |

Table - Masses distribution (computed using T1100261)

Acceptance Criteria

The Mass budget must be

- 579.1 Kg (cf E1100427)+/-25Kg (5%)

Test result:	Passed:   X   	Failed:         .

## Step 8: Lockers adjustment

| D.I at Lockers | Dial indicators V | Dial indicators H |
| --- | --- | --- |
| A | 0 | 1 |
| B | 1.5 | -2 |
| C | 0 | 0.5 |
| D | -1 | -1 |

Table – Dial indicators read-out (in thousands of an inch)

Acceptance Criteria

- Vertical and horizontal displacement near the lockers must be lower than 2 mils (0.002”)

Test result:	Passed:   X    	Failed:         .

## Step 9 - Electronics Inventory

| Hardware | LIGO reference | S/N |
| --- | --- | --- |
| Coil driver | D0902744 | S1103306 |
|  |  | S1103304 |
| Anti Image filter | D070081 | S1104573 |
| Anti aliasing filter | D1000269 | S1104632 |
|  |  | S104635 |
| Interface chassis | D1000067 | S1107436 |
|  |  | S1107437 |

Table - Inventory electronics

Acceptance Criteria

- Inventory is complete

Test result:	Passed:   X   	Failed:         .

## Step 10 – Cables inventory – E1100822

The location of all cables must be reported in the spreadsheet E1100822.

| Cable Connects | Cable S/N |
| --- | --- |
| Part Name | Configuration | Corner 1 | Corner 2 | Corner 3 |
| GS13 | Horizontal | S1104692-S1104657 | S1104607-S1104648 | S1104597-S110650 |
| GS13 | Vertical |  |  |  |
| L4C | Horizontal | N/A | N/A | N/A |
| L4C | Vertical | N/A | N/A | N/A |
| Actuator | Horizontal | S1106680 | S1104748 | S1104718 |
|  | Vertical | S1104749 | S1104498 | S1104483 |

Table – Cables inventory

Acceptance Criteria

- Cable inventory completed

- E110082 spreadsheet updated

Test result:	Passed:   X   	Failed:         .

## Step 11 - Set up sensors gap

|  | 10 Kg masses at each corners | No mass |
| --- | --- | --- |
| Table locked | ADE boxes on | ADE boxes on |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Std deviation |
| H1 |  |  | 190 |  |
| H2 |  |  | -370 |  |
| H3 |  |  | 5 |  |
| V1 |  |  | -490 |  |
| V2 |  |  | 460 |  |
| V3 |  |  | -180 |  |

Table – Capacitive position sensor readout after gap set-up

Issues/difficulties/comments regarding this test: HAM-ISI – LLO HAM 3 uses synchronized boxes, with power boards installed on their back.

Note that the position sensors were originally set up this way but that later adjustments to the lockers brought them to the level shown in Step 5, then judged acceptable.

Acceptance criteria:

- All mean values must be lower than 400 cts (a bit less than .0005”).

- All standard deviations below 5 counts.

- No cross talk

Test result:	Passed: X 	Failed:       .

## Step 12 - Measure the Sensor gap

| Sensors | Gap measured on the Jig | Gap measured on the table |
| --- | --- | --- |
| H1 | NR | 80 |
| H2 | NR | 80 |
| H3 | NR | 85 |
| V1 | NR | 80 |
| V2 | NR | 85 |
| V3 | NR | 85 |

Acceptance criteria:

Sensors gap measured on the jig and on the optic table must be:

- 0.080” +/-0.002”

Test result:	Passed:     	Failed:    X     .

## Step 13 - Check Sensor gaps after the platform release

|  | Table locked | Table unlocked |
| --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Difference |
| H1 | -2230.9 |  | -2086.9 | 144 |
| H2 | -636.48 |  | -66.072 | 570.408 |
| H3 | -1300.2 |  | -459.22 | 840.98 |
| V1 | -728.31 |  | -607.66 | 120.65 |
| V2 | 797.89 |  | 825.84 | 27.95 |
| `V3 | -715.5 |  | -1386.5 | -671 |

Table – Sensor gaps after platform release

Acceptance criteria:

- Absolute values of the difference between the unlocked and the locked table must be below:

  - 1600 cts for horizontal sensors (~0.002”)

  - 1600 cts for vertical sensors (~0.002”)

- Considering the acceptance criteria of step 4, all mean values must be lower than

  - 2000 cts for horizontal sensors (~0.0025”)

  - 2000 cts for vertical sensors (~0.0025”)

Test result:	Passed:  X    	Failed:         .

## Step 14– Performance of the limiter

## Step 14.1 - Test Nº1 - Push “in the general coordinates”

We conducted this test because locker C would not unlock all the way. We believe that the table was completely free but ran that test in order to measure whether this was true or not.

Numbers could not recorded because of a workstation issue, but in all instances, they were > 18,000 counts. (which is superior to the 16,000 counts criteria)

Also, the direction were right.

Test result:	Passed:   X   	Failed:         .

## Step 15 - Position Sensors unlocked/locked  Amplitude Spectral Densities

As scheduled, we installed new horizontal GS-13 seismometers on the ISI. As the table was still locked (because weight repartition had not been finalized), we took Amplitude Spectral Density measurements and noticed that the H3 sensor did not have a normal response.

We switched the pod used in H3.

Figure 1 Calibrated GS-13 ASD plot- Broken Sensor behavior

As we can see in the top plot, the seismometer had a normal behavior down to ~13 Hz, had a somewhat low answer between 2 and 13 Hz, and was downright not working at any lower frequencies.

On the lower plot, we can see the ASD when a pod is disconnected (which happened on V1 as we were working on switching the H3 pod). So, the cyan line in this plot displays the ASD of the GS-13 electronics (from the feedthrus up), and the calibration.

Data files in SVN at:

/seismic/HAM-ISI/L1/HAM3/Data/Spectra/Undamped/

- LLO_ISI_HAM3_ASD_m_CPS_T240_L4C_GS13_Locked_vs_Unlocked_2012_01_31.pdf

Scripts files for taking and processing the data, and plotting it in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Plot_ASD_Unlocked_Locked_HAM_ISI.m

- Plot_ASD_Unlocked_Locked_Group_HAM_ISI.m

Figures in SVN at:

seismic/HAM-ISI/L1/HAM3/Data/Figures/Spectra/Undamped/

- LLO_ISI_HAM3_ASD_m_CPS_Requirements_Locked_vs_Unlocked_2012_01_31.pdf

- LLO_ISI_HAM3_ASD_m_GS13_Requirements_Locked_vs_Unlocked_2012_01_31.pdf

CPS calibration:

The CPS power spectrums are calibrated by using a sensitivity of 30.2 nm/count.

Figure - Calibrated CPS power spectrum

Figure – Power spectrum Calibrated GS13

Acceptance criteria:

- No cross talk (peaks at low frequencies + harmonics on measurements)

- Magnitudes of power spectra must be between requirement curves such as in the following figures (dashed lines)

Test result:	Passed:   X   	Failed:          .

## Step 16 - GS13 ASD -tabled tilted

The figure below presents the GS13 power spectrum when the table is unlocked and loaded with a 20Kg mass at one of its corner.

Figure – ASD Calibrated GS13 with mass at corner

Data files in SVN at:

/seismic/HAM-ISI/L1/HAM3/Data/Spectra/Undamped/

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_GS13_Table_Tilted_2011-03-31.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Plot_ASD_Tilted_Stage_HAM_ISI.m

Figures in SVN at:

seismic/HAM-ISI/L1/HAM3/Data/Figures/Spectra/Undamped/

- LLO_ISI_HAM3_ASD_CT_GS13_Tilted_2012_01_31.pdf

Acceptance criteria:

- With table unlocked and tilted, magnitudes of power spectra must be fully included within:

Issues/difficulties/comments regarding this test:

As can be seen on the plot, in 4/6 tilted table configurations, the H3 seismometer does not follow the same trend around the 1Hz resonance (of the system). We confirmed those results on 02/02/12 by taking similar measurements.

Then, to ensure that this was describing actual motion and a sensor issue, we plotted the corresponding graphs for the CPS sensors:

Figure – ASD Calibrated CPS with mass at corner

Those confirmed that we had a strong motion reduction on H3  with a mass added on locations B, C, E and F, confirming our suspicions that the H3 inertial sensor was working fine.

We also ran 2 additional tests, slowly adding mass on corners E and F, and taking ASD at each additional mass added. These 2 tests additionally confirmed that at some level when adding the mass, the table got locked in a position in which the H3 sensors would not see any motion.

Test result:	Passed:   X   	Failed:          .

## Step 17- GS13 pressure readout

All vertical seismometers were reading 77 or 78 kPa, whereas the differential readout was consistently reading 0.

Test result:	Passed:   X   	Failed:       .

## Step 18 - Actuators Sign and range of motion (Local drive)

|  | Negative drive | Initial offset | Positive drive |
| --- | --- | --- | --- |
| H1 readout (count) | 15600 | -2086.6 | -23674 |
| H2 readout (count) | 22696 | -66.07 | -23804 |
| H3 readout (count) | 22721 | -459.22 | -26017 |
| V1 readout (count) | 18674 | -607.66 | -20239 |
| V2 readout (count) | 26334 | 825.84 | -23952 |
| V3 readout (count) | 21736 | -1386.5 | -21546 |

Table - Range of motion - Local drive

Data files in SVN at:

/seismic/HAM-ISI/L1/HAM3/Data/Static_Tests/

- LLO_ISI_HAM3_Range_Of_Motion_20120131.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Range_Motion_HAM_ISI.m

Acceptance criteria:

- Main couplings sensors readout must be at least 16000 counts (~0.02”)

- A positive offset drive on one actuator must give positive sensor readout on the collocated sensor. Signs will also be tested when measuring local to local transfer functions.

Test result:	Passed:   X   	Failed:         .

## Step 19 - Static Testing (Tests in the local basis)

|  | H1 | H2 | H3 | V1 | V2 | V3 |
| --- | --- | --- | --- | --- | --- | --- |
| H1 | -1882.9 | -1176.2 | -1170.8 | 0.06016 | 21.654 | 7.3542 |
| H2 | -1202.1 | -1914.3 | -1199.1 | 5.6833 | 9.6504 | 10.663 |
| H3 | -1160.7 | -1163.2 | -1868.4 | 13.998 | -26.884 | 6.3794 |
| V1 | -181.36 | -177.29 | 341.02 | -1351.9 | 54.534 | 578.36 |
| V2 | 339.77 | -173.64 | -170.13 | 560.06 | -1308.2 | 48.246 |
| V3 | -169.88 | 357.92 | -166.17 | 16.386 | 592.07 | -1347.9 |

Table - Main and cross coupling

Data files in SVN at:

/seismic/HAM-ISI/L1/HAM3/Data/Static_Tests/

- LLO_ISI_HAM3_Offset_Local_Drive_20120130.mat

Scripts files for taking data in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Static_Test_Local_Basis_HAM_ISI.m

Acceptance criteria:

- Vertical

For a +1000 count offset drive on vertical actuators

  - Collocated sensors must be 1400 counts +/- 10%

- Horizontal

For a +1000 count offset drive on horizontal actuators

  - Collocated sensors must be 2000 counts +/- 10%

  - Non-collocated horizontal sensors must be 1250 counts +/-10%

Test result:	Passed:   X    	Failed:         .

## Step 20 - Linearity test

|  | Slope | Offset | Average slope | Variation from average(%) |
| --- | --- | --- | --- | --- |
| H1 | -1.877 | -2115.5 | -1.8847 | -0.4068 |  |
| H2 | -1.9143 | -85.46 |  | 1.5723 |  |
| H3 | -1.8627 | -469.99 |  | -1.1655 |  |
| V1 | -1.3472 | -590.08 | -1.3396 | 0.5648 |  |
| V2 | -1.3273 | 843.02 |  | -0.9206 |  |
| V3 | -1.3444 | -1365.5 |  | 0.3558 |  |

Table - Slopes and offset of the triplet Actuators - HAM-ISI - Sensors

Scripts files for taking data in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Scripts/Data_Collection

- Linearity_Test_Awgstream_HAM_ISI.m

Figure - Horizontal and vertical actuators x HAM-ISI x sensors

Data files in SVN at:

seismic/HAM-ISI/L1/HAM3/Data/Linearity_Test/

- LLO_ISI_HAM3_Linearity_test_20120131.mat

Figures in SVN at:

seismic/HAM-ISI/L1/HAM3/Data/Figures/Linearity_Test/

- LLO_ISI_HAM3_Linearity_test_20120131.pdf

Acceptance criteria:

- Horizontal and vertical slopes of the triplet actuators x HAM-ISI x sensors:  Average slope +/- 1%

Test result:	Passed:      	Failed:   X

## Step 21 Cartesian Basis Static Testing

Table - Tests in the general coordinate basis

Issues/difficulties/comments regarding this test:

This is the first ISI oriented in the Y direction. Matrices were redesigned for that purpose.

Acceptance criteria:

For a positive drive in the Cartesian basis:

- Local sensor readout must have the same sign that the reference table (CONT2ACT check)

- Cartesian sensors read out must be positive (DISP2CEN check) in the drive direction

Test result:	Passed:      	Failed:         .

## Step 22- Frequency response

Compensation filters of the new GS13 interface chassis are located in the geophone pre-filters bank.

## Step 22.1 - Local to local measurements

Local to local transfer functions have been measured with 90 repetitions.

Data files in SVN at:

seismic/HAM-ISI/L1/HAM3/Data/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_HAM3_Data_TF_L2L_50mHz_500mHz_20120201-000859.mat

- LLO_ISI_HAM3_Data_TF_L2L_500mHz_5Hz_20120131-213926.mat

- LLO_ISI_HAM3_Data_TF_L2L_5Hz_200Hz_20120131-205053.mat

- LLO_ISI_HAM3_Data_TF_L2L_200Hz_800Hz_20120131-191720.mat

Data collection script files:

seismic/HAM-ISI/L1/HAM3/Scripts/Data_Collection/

- Run_Exc_Batch_L1ISIHAM3.m

Scripts files for processing and plotting in SVN at:

seismic/HAM-ISI/L1/HAM3/Scripts/Control_Scripts/

- Step_1_TF_L2L_L1_ISI_HAM3.m

Figures in SVN at:

/seismic/HAM-ISI/L1/HAM3/Data/Figures/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_HAM3_TF_L2L_Raw_from_ACT_to_CPS_2012_01_31.fig

- LLO_ISI_HAM3_TF_L2L_Raw_from_ACT_to_GS13_2012_01_31.fig

Storage of measured transfer functions in the SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_functions/ Measurements/

Undamped/

- LLO_HAM_ISI_Unit_3_Data_TF_L2L_2011_04_06.mat

The local to local transfer functions are presented below.

Figure - Local to Local Measurements –Capacitive Position Sensors

Figure - Local to Local Measurements – Inertial sensors

Issues/difficulties/comments regarding this test:

Around 1 Hz, the H2 inertial sensor seems to indicate different behaviors of each corner.

## Step 23 - Lower Zero Moment Plane

Based from results of Phase I where the measurement was taken with, we wanted to retake this measurement. The measurements were run but data processing issues have so far prevented us from getting the results.

Test result:	Passed:   X    	Failed:         .

Conclusion of Side Chamber testing

Very few issues were found during the testing of this unit. The known issues are summed up here:

- The Linearity test gives a slightly different slope on the H2 actuator (1.5% from average for a 1% criteria)

- The H3 GS-13 gives a slightly different response around 1 Hz when the table is tilted.

- We have a slight offset of 2 CPS: H1 and V3.

- The locked table measurements seem to indicate that the table was not completely locked.

Issues unchanged since the Phase I testing:

- sensor gaps not recorded on the jig

- Vertical spring constant: one spring appears to have a much lower constant than the others (-8.91% than the average of the 3).

- LZMP measurements were done with only 1 average (high uncertainty), they should be redone inside the chamber

# INITIAL IN CHAMBER TESTING

The HAM-ISI was installed in HAM 3 on February 3rd 2012. This initial in chamber testing was conducted in the following 2 weeks until we closed the chamber on February 17th 2012.

## Step 1: Check level of Stage 1 Optical Table

Issues/difficulties/comments regarding this test:

Initially this test was failed, but we adjusted the shim thickness to pass this test (see following step for new locker shims).

The table here does not appear to be really tilted. Most of the difference can be included within the .005” flatness of the table.

Acceptance Criteria

- The maximum angle of the table with the horizontal mustn’t exceed  ~100µrad

Test result:	Passed:   X   	Failed:         .

## Step 2: Shim thickness

Issues/difficulties/comments regarding this test: Those shims were adjusted -in comparison to side chamber testing to meet level requirements. We believe it was due to 2 things: our level slightly out calibration and our test stand off.

| Lockers | Shim thickness (mil) |
| --- | --- |
| A | 130 |
| B | 125 |
| C | 120 |
| D | 125 |

Table – Shims Thickness

Acceptance Criteria

- Inventory is complete

Test result:	Passed:   X   	Failed:         .

## Step 5: Blade spring profile

Because we adjusted the shims, we retook this measurement:

| Blade # | Base (") | Tip(") | Flatness (mils) |
| --- | --- | --- | --- |
| 1 | .4961 | .4890 | +7.1 |
| 2 | .4986 | .4914 | +7.2 |
| 3 | .4978 | .4882 | +9.6 |

Table 2 - Blade profile

Acceptance Criteria:

- Blades must be flat within 0.015" inches.

Test result:	Passed:   X   	Failed:         .

## Step 7: Mass budget

Figure – Keel Masses and Wall masses location

Figure  - Optical table masses distribution

|  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| w9 | 1 | 1 |  |  |  | 1 |  | 17.3 | 7.85 |
| w1 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w2 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w3 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| w4 |  | 2 |  | 1 |  | 1 | 1 | 49.5 | 22.45 |
| w5 | 1 | 2 |  |  |  | 1 | 1 | 45.6 | 20.68 |
| w6 |  | 3 |  |  |  | 1 | 1 | 46.1 | 20.91 |
| w7 |  | 3 |  |  |  | 1 | 1 | 46.1 | 20.91 |
| w8 |  |  |  |  |  | 1 | 1 | 42.8 | 19.41 |
| Side Masses Total | 2 | 11 | 0 | 1 | 0 | 9 | 8 | 375.8 | 170.46 |

Table – Wall masses distribution

|  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| k1 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k2 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
| k3 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k4 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
| k5 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k6 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
|  | 3 | 3 | 3 | 0 | 3 | 0 | 6 | 198.6 | 90.08 |

Table – Keel masses distribution

|  | D972213 | D972215 | D0901075 |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | 2.5 kg | 5 kg | 10 kg |  |  |
|  | 610 | 230 | 5.5 | 11 | 22 | lbs | kgs |
| A | 1 |  |  | 1 |  | 621 | 281.68 |
| B |  |  |  | 1 |  | 11 | 4.99 |
| C |  |  |  | 1 |  | 11 | 4.99 |
| D |  |  |  | 1 |  | 11 | 4.99 |
| E-1 |  |  |  | 1 |  | 11 | 4.99 |
| E-2 |  |  |  | 1 |  | 11 | 4.99 |
| E-3 |  |  |  | 1 |  | 11 | 4.99 |
| Top Masses | 1 | 0 | 0 | 7 | 0 | 687 | 311.62 |

Table - Optical Table Masses distribution

|  | Side | Keel | Top | Total |
| --- | --- | --- | --- | --- |
| Weigh (kg) | 170.46 | 90.08 | 311.62 | 572.16 |

Table - Masses distribution (computed using T1100261)

Acceptance Criteria

The Mass budget must be

- 579.1 Kg (cf E1100427)+/-25Kg (5%)

Test result:	Passed:   X   	Failed:         .

## Step 8: Lockers adjustment

- Not recorded but within.

Test result:	Passed:   X    	Failed:         .

.

## Step 10 – Cables inventory – E1100822

We noticed a few mistakes in our early installation of cables, so here is a corrected inventory

| Cable Connects | Cable S/N |
| --- | --- |
| Part Name | Configuration | Corner 1 | Corner 2 | Corner 3 |
| GS13 | Horizontal | S1104692-S1104657 | S1104607-S1104648 | S1104597-S110650 |
| GS13 | Vertical |  |  |  |
| L4C | Horizontal | N/A | N/A | N/A |
| L4C | Vertical | N/A | N/A | N/A |
| Actuator | Horizontal | S1106680 | S1104748 | S1104486 |
|  | Vertical | S1104749 | S1104718 | S1104483 |

Table – Cables inventory

Acceptance Criteria

- Cable inventory completed

- E110082 spreadsheet updated

Test result:	Passed:   X   	Failed:         .

## Step 11 - Set up sensors gap

|  | No mass |
| --- | --- |
| Table locked | ADE boxes on |
| Sensors | Offset (Mean) | Std deviation |
| H1 | 190 |  |
| H2 | -370 |  |
| H3 | 5 |  |
| V1 | -490 |  |
| V2 | 460 |  |
| V3 | -180 |  |

Table – Capacitive position sensor readout after gap set-up

Issues/difficulties/comments regarding this test: HAM-ISI – LLO HAM 3 uses synchronized boxes, with power boards installed on their back.

Note that the position sensors were originally set up this way but that later adjustments to the lockers brought them to the level shown in Step 5, then judged acceptable.

Acceptance criteria:

- All mean values must be lower than 400 cts (a bit less than .0005”).

- All standard deviations below 5 counts.

- No cross talk

Test result:	Passed: X 	Failed:       .

## Step 13 - Check Sensor gaps after the platform release

|  | Table locked | Table unlocked |
| --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Difference |
| H1 | -2230.9 |  | -2086.9 | 144 |
| H2 | -636.48 |  | -66.072 | 570.408 |
| H3 | -1300.2 |  | -459.22 | 840.98 |
| V1 | -728.31 |  | -607.66 | 120.65 |
| V2 | 797.89 |  | 825.84 | 27.95 |
| `V3 | -715.5 |  | -1386.5 | -671 |

Table – Sensor gaps after platform release

Acceptance criteria:

- Absolute values of the difference between the unlocked and the locked table must be below:

  - 1600 cts for horizontal sensors (~0.002”)

  - 1600 cts for vertical sensors (~0.002”)

- Considering the acceptance criteria of step 4, all mean values must be lower than

  - 2000 cts for horizontal sensors (~0.0025”)

  - 2000 cts for vertical sensors (~0.0025”)

Test result:	Passed:  X    	Failed:         .

## Step 15 - Position Sensors unlocked/locked  Amplitude Spectral Densities

Data files in SVN at:

/seismic/HAM-ISI/L1/HAM3/Data/Spectra/Undamped/

- LLO_ISI_HAM3_ASD_m_CPS_T240_L4C_GS13_Locked_vs_Unlocked_2012_02_16.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Plot_ASD_Unlocked_Locked_HAM_ISI.m

- Plot_ASD_Unlocked_Locked_Group_HAM_ISI.m

Figures in SVN at:

seismic/HAM-ISI/L1/HAM3/Data/Figures/Spectra/Undamped/

- LLO_ISI_HAM3_ASD_m_CPS_Requirements_Locked_vs_Unlocked_2012_02_16.pdf

- LLO_ISI_HAM3_ASD_m_GS13_Requirements_Locked_vs_Unlocked_2012_02_16.pdf

CPS calibration:

The CPS power spectrums are calibrated by using a sensitivity of 30.2 nm/count.

Figure - Calibrated CPS power spectrum

Figure – Power spectrum Calibrated GS13

Acceptance criteria:

- No cross talk (peaks at low frequencies + harmonics on measurements)

- Magnitudes of power spectra must be between requirement curves such as in the following figures (dashed lines)

Test result:	Passed:   X   	Failed:          .

## Step 16 - GS13 ASD -tabled tilted

The figure below presents the GS13 power spectrum when the table is unlocked and loaded with a 20Kg mass at one of its corner.

Figure – ASD Calibrated GS13 with mass at corner

Data files in SVN at:

/seismic/HAM-ISI/L1/HAM3/Data/Spectra/Undamped/

- LLO_HAM_ISI_Unit_3_Calibrated_PSD_GS13_Table_Tilted_2011-06_16.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Plot_ASD_Tilted_Stage_HAM_ISI.m

Figures in SVN at: seismic/HAM-ISI/L1/HAM3/Data/Figures/Spectra/Undamped/

- LLO_ISI_HAM3_ASD_CT_GS13_Tilted_2012_02_16.pdf

Issues/difficulties/comments regarding this test:

The issue that we had during side chamber testing was found to be combination of 2 factors:

- 1 peek cable clamp was  too close to Stage 1 and making intermittent contact

- we were not using enough weight and used 20 lbs instead of 40.

Test result:	Passed:   X   	Failed:          .

## Step 17- GS13 pressure readout

All vertical seismometers were reading 77 or 78 kPa, whereas the differential readout was consistently reading 0.

Test result:	Passed:   X   	Failed:       .

## Step 18 - Actuators Sign and range of motion (Local drive)

|  | Negative drive | Positive drive |
| --- | --- | --- |
| H1 readout (count) | -21614 | 26101 |
| H2 readout (count) | -22249 | 25593 |
| H3 readout (count) | -24183 | 24696 |
| V1 readout (count) | -20683 | 18431 |
| V2 readout (count) | -28251 | 22259 |
| V3 readout (count) | -21034 | 22066 |

Table - Range of motion - Local drive

Data files in SVN at:

/seismic/HAM-ISI/L1/HAM3/Data/Static_Tests/

- LLO_ISI_HAM3_Range_Of_Motion_20120215.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Range_Motion_HAM_ISI.m

Issues/difficulties/comments regarding this test:

Signs are opposed to what we found side-chamber. This is due to the fact that we have here the compensation filters engaged  in the actuators filter banks, as they should be.

Acceptance criteria:

- Main couplings sensors readout must be at least 16000 counts (~0.02”)

- A positive offset drive on one actuator must give positive sensor readout on the collocated sensor. Signs will also be tested when measuring local to local transfer functions.

Test result:	Passed:   X   	Failed:         .

## Step 20 - Linearity test

|  | Slope | Offset | Average slope | Variation from average(%) |
| --- | --- | --- | --- | --- |
| H1 | 1.8600 | 2009.8 | 1.86276667 | 0.148524596 |  |
| H2 | 1.8913 | 928.2430 |  | -1.53177174 |  |
| H3 | 1.8370 | -1129.2 |  | 1.383247141 |  |
| V1 | 1.3427 | -2401.5 | 1.3388 | -0.29130565 |  |
| V2 | 1.3393 | -26397 |  | -0.03734688 |  |
| V3 | 1.3344 | 3558.8 |  | 0.328652525 |  |

Table - Slopes and offset of the triplet Actuators - HAM-ISI - Sensors

Scripts files for taking data in SVN at: seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Linearity_Test_Awgstream_HAM_ISI.m

Figure - Horizontal and vertical actuators x HAM-ISI x sensors

Data files in SVN at:  seismic/HAM-ISI/L1/HAM3/Data/Linearity_Test/

- LLO_ISI_HAM3_Linearity_test_20120215.mat

Figures in SVN at:

seismic/HAM-ISI/L1/HAM3/Data/Figures/Linearity_Test/

- LLO_ISI_HAM3_Linearity_test_20120215.pdf

Issues/difficulties/comments regarding this test:

One can notice that the slopes are opposed to the ones in the same test side-chamber. This is due to the fact that we have here the compensation filters engaged  in the actuators filter banks as they should be.

Acceptance criteria:

- Horizontal and vertical slopes of the triplet actuators x HAM-ISI x sensors:  Average slope +/- 1%

Test result:	Passed:      	Failed:   X

## Step 22- Frequency response

Compensation filters of the new GS13 interface chassis are located in the geophone pre-filters bank.

## Step 22.1 - Local to local measurements

Local to local transfer functions have been measured with 90 repetitions, and those were taken once the door of the chamber had been closed..

Data files in SVN at:

seismic/HAM-ISI/L1/HAM3/Data/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_HAM3_Data_TF_L2L_50mHz_500mHz_20120201-000859.mat

- LLO_ISI_HAM3_Data_TF_L2L_500mHz_5Hz_20120131-213926.mat

- LLO_ISI_HAM3_Data_TF_L2L_5Hz_200Hz_20120131-205053.mat

- LLO_ISI_HAM3_Data_TF_L2L_200Hz_800Hz_20120131-191720.mat

Data collection script files:

seismic/HAM-ISI/L1/HAM3/Scripts/Data_Collection/

- Run_Exc_Batch_L1ISIHAM3.m

Scripts files for processing and plotting in SVN at:

seismic/HAM-ISI/L1/HAM3/Scripts/Control_Scripts/

- Step_1_TF_L2L_L1_ISI_HAM3.m

Figures in SVN at:

/seismic/HAM-ISI/L1/HAM3/Data/Figures/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_HAM3_TF_L2L_Raw_from_ACT_to_CPS_2012_02_18.fig

- LLO_ISI_HAM3_TF_L2L_Raw_from_ACT_to_GS13_2012_02_18.fig

Storage of measured transfer functions in the SVN at:

opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_functions/ Measurements/

Undamped/

- LLO_HAM_ISI_Unit_3_Data_TF_L2L_2012.mat

The local to local transfer functions are presented below.

Figure - Local to Local Measurements –Capacitive Position Sensors

Figure - Local to Local Measurements – Inertial sensors

## Step 23 - Lower Zero Moment Plane

Data collection script files: seismic/HAM-ISI/Common/Transfer_Functions_Scripts

- Run_Cart2Cart_10mHz_100mHz.m

Data files in SVN at: seismic/HAM-ISI/L1/HAM3/Data/Transfer_Functions/Measurements/

Undamped

- LLO_ISI_HAM3_Data_TF_C2C_10mHz_100mHz_LZMP_20120218-185146.mat

Scripts files for processing and plotting in SVN at:

/seismic/HAM-ISI/X2/Data/Unit_3/Transfer_Functions/Measurements/Undamped/

- LZMP_2012.m

Figures in SVN at : seismic/HAM-ISI/L1/HAM3/Data/Figures/Transfer_Functions/

Measurements/Undamped

- LLO_ISI_HAM3_LZMP_20120218.fig

X & Y offsets:

| X offset (mm) | 1.2024 |
| --- | --- |
| Y offset (mm) | 0.22107 |

Table – Offset of the Lower Zero Moment Plane

The results from two measurements are presented on the figure below:

Figure - Lower Zero moment plane – Main and cross couplings at low frequency

Issues/difficulties/comments regarding this test

This measurement was done

Acceptance criteria:

- X offset must be less than 2 mm

- Y offset must be less than 2 mm

Test result:	Passed:   X    	Failed:         .

Conclusion of Initial In-Chamber testing

Most issues found while side-testing were solved during this phase of the testing. Very few issues were found during the testing of this unit. The known issues are summed up here:

- The Linearity test gives a slightly different slope on the H2 actuator (1.5% from average for a 1% criteria)

- We have a slight offset of 2 CPS: H1 and V3.

Issues unchanged since the Phase I testing:

- sensor gaps not recorded on the jig

- Vertical spring constant: one spring appears to have a much lower constant than the others (-8.91% than the average of the 3).