LIGO Laboratory / LIGO Scientific Collaboration

LIGO- E1200105 	                 LIGO                   March 15th2012

aLIGO HAM-ISI, Installation Test Report, Phase II,

LLO HAM-ISI 2

E1200105-v2

Adrien Le Roux, Joe Hanson, Celine Ramet, Michael Vargas

Distribution of this document:

Advanced LIGO Project

This is an internal working note

of the LIGO Laboratory

| California Institute of Technology LIGO Project – MS 18-34 1200 E. California Blvd. Pasadena, CA 91125 Phone (626) 395-2129 Fax (626) 304-9834 E-mail: info@ligo.caltech.edu | Massachusetts Institute of Technology LIGO Project – NW22-295 185 Albany St Cambridge, MA 02139 Phone (617) 253-4824 Fax (617) 253-7014 E-mail: info@ligo.mit.edu |
| --- | --- |
| LIGO Hanford Observatory P.O. Box 1970 Mail Stop S9-02 Richland WA 99352 Phone 509-372-8106 Fax 509-372-8137 | LIGO Livingston Observatory P.O. Box 940 Livingston, LA 70754 Phone 225-686-3100 Fax 225-686-7189 |

Table of contents:

Introduction	3

I.	SIDE CHAMBER TESTING	4

	Step 1: GS13	4

	Step 1.1 – Horizontal GS-13s	4

	Step 2: Inventory (E1000052)	6

	Step 3: Blade spring profile	7

	Step 4: Gap checks on actuators-after installation on Stage 1	7

	Step 5: Check level of Stage 1 Optical Table	8

	Step 6: Mass budget	9

	Step 7: Shim thickness	10

	Step 8: Lockers adjustment	10

	Step 9 - Electronics Inventory	11

	Step 10 – Cables inventory – E1100822	11

	Step 10 - Set up sensors gap	11

	Step 11 - Check Sensor gaps after the platform release	12

	Step 12 - Position Sensors unlocked/locked Power Spectrum	12

	Step 13- GS13 power spectrum -tabled tilted	15

	Step 14- GS13 pressure readout	16

	Step 15- Actuators Sign and range of motion (Local drive)	16

	Step 16 - Static Testing (Tests in the local basis)	17

	Step 17 - Linearity test	17

	Step 18 - Cartesian Basis Static Testing	19

	Step 19- Frequency response	19

	Step 19.1 - Local to local measurements	19

	Step 20 - Lower Zero Moment Plane	24

Conclusion Side Chamber testing	25

II.	INITIAL CHAMBER TESTING	26

	Step 5: Check level of Stage 1 Optical Table	26

	Step 3: Blade spring profile	27

	Step 4: Gap checks on actuators-after installation on Stage 1	27

	Step 6: Mass budget	27

	Step 7: Shim thickness	29

	Step 8: Lockers adjustment	29

	Step 9 - Electronics Inventory	29

	Step 10 - Set up sensors gap	30

	Step 11 - Check Sensor gaps after the platform release	30

	Step 12 - Position Sensors unlocked/locked Power Spectrum	31

	Step 13- GS13 power spectrum -tabled tilted	33

	Step 14- GS13 pressure readout	34

	Step 15- Actuators Sign and range of motion (Local drive)	34

	Step 16 - Static Testing (Tests in the local basis)	35

	Step 17 - Linearity test	35

	Step 19- Frequency response	37

	Step 19.1 - Local to local measurements	37

Conclusion Initial In-Chamber testing	38

Introduction

This unit was first assembled in the early fall of 2010, and tested until October 2010. Following the discovery of un-autorized repairs while building this unit, the testing was interrupted. Testing done until interruption can be seen under v1 of this document. The unit was subsequently disassembled in Spring 2011 and immediately re-assembled.

The unit was pulled from storage in February 2012, and the assembly was then completed (3 new horizontal GS-13s, Tuned Mass Dampers and final cabling were installed).

Testing was completed March 3rd 2012.

# SIDE CHAMBER TESTING

## Step 1: GS13

All data related to GS-13 post podding testing can be found in the SVN at : SeismicSVN\seismic\Common\Data\aLIGO_GS13_TestData\PostMod_TestResults_PDFs.

E1000058 spreadsheet provides the status of each individual GS-13 at LLO site during aLIGO HAM assembly

Data files in SVN at:

/opt/svncommon/seisvn/seismic/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_RawASCII

Scripts files for processing and plotting in SVN at: seismic/Common/MatlabTools

- gs13qatest.m

Figures in SVN at:

/opt/svncommon/seisvn/seismic/Common/Data/aLIGO_GS13_TestData/PostMod_TestResults_PDFs

## Step 1.1 – Horizontal GS-13s

Huddle testing

Figure - Huddle testing of Horiz GS-13 809, 787 and 706 after aLIGO modifications

We can notice that both 787 and 776 appear to have their resonances slightly off from 1Hz.

Acceptance Criteria:

- GS13 should have been already modified and tested. GS-13 Inspection/Pod Assembly (D047810). Checklist is defined in F090070-v6

- Resonant frequency at 1 Hz (huddle testing)

- No spring resonance on vertical driven tests

Test result:	Passed:   X  	Failed:         .

## Step 2: Inventory (E1000052)

| DCC/Vendor number | Part name | Configuration | S/N | S/N | S/N |
| --- | --- | --- | --- | --- | --- |
| D071001 | Stage 0 base |  | 002 |  |  |
| D071050 | Stage 1 base |  | 001 |  |  |
| D071051 | Optical table |  | 002 |  |  |
| D071002 | Spring Post |  | 007 | 005 | 015 |
| D071100 | Spring |  | 34 | 17 | 39 |
| D071102 | Flexure |  | 42 | 21 | 45 |
| ADE | Position sensor | Horizontal | 11981 | 12079 | 12071 |
|  |  | Vertical | 12078 | 12080 | 12069 |
| D047812 | GS-13 pod | Horizontal | 46 | 39 | 73 |
|  |  | Vertical | 32 | 19 | 51 |
| D047823 | L4C pod | Horizontal | N/A | N/A | N/A |
|  |  | Vertical | N/A | N/A | N/A |
| D0902749 | Actuator | Horizontal | L139 | L081 | L086 |
|  |  | Vertical | L084 | L089 | L092 |

## Step 3: Blade spring profile

Re-measured because we had to change shims to meet the platform level requirements.

| Blade # | Base (") | Tip(") | Flatness (mils) |
| --- | --- | --- | --- |
| 1 | .484 | .473 | +11 |
| 2 | .486 | .470 | +16 |
| 3 | .484 | .473 | +11 |

Table 1 - Blade profile

Issues/difficulties/comments regarding this test:

Those were re-measured because we adjusted the shims to get the table within level requirements.

Acceptance Criteria:

- Blades must be flat within 0.015" inches.

Test result:	Passed:      	Failed:  X       .

## Step 4: Gap checks on actuators-after installation on Stage 1

Not recorded but checked and within requirements.

Acceptance Criteria

- Gaps must be within 0.010” of design value (0.080”)

Test result:	Passed:   X  	Failed:         .

## Step 5: Check level of Stage 1 Optical Table

Acceptance Criteria

- The maximum angle of the table with the horizontal mustn’t exceed  ~100µrad

This table appears to have local deformation but no excessive tilt

Test result:	Passed:   X  	Failed:         .

## Step 6: Mass budget

Figure – Keel Masses and Wall masses location

|  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| k1 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k2 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
| k3 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k4 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
| k5 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k6 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
|  | 3 | 3 | 3 | 0 | 3 | 0 | 6 | 198.6 | 90.08 |

Table- Keel masses

|  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| w9 | 1 | 1 | 1 | 1 |  | 1 |  | 24 | 10.89 |
| w1 | 2 | 1 | 1 | 1 |  | 1 |  | 24.6 | 11.16 |
| w2 | 3 |  | 2 | 1 |  | 1 |  | 26.3 | 11.93 |
| w3 | 2 |  | 1 | 1 |  | 1 |  | 23.5 | 10.66 |
| w4 | 3 |  | 1 | 1 |  | 1 |  | 24.1 | 10.93 |
| w5 |  |  | 1 | 1 |  | 1 |  | 22.3 | 10.12 |
| w6 | 2 | 1 | 1 | 1 |  | 1 |  | 24.6 | 11.16 |
| w7 | 1 |  |  | 1 |  | 1 |  | 20.7 | 9.39 |
| w8 | 3 |  | 2 | 1 |  | 1 |  | 26.3 | 11.93 |
| Side Masses Total | 17 | 3 | 10 | 9 | 0 | 9 | 0 | 216.4 | 98.16 |

Table- Side masses

|  | D972213 | D972214 | D972215 | D0901075 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | 2.5 kg | 5 kg | 10 kg |  |  |
|  | 610 | 375 | 230 | 5.5 | 11 | 22 | lbs | kgs |
| Top Masses | 0 | 2 | 0 | 0 | 3 | 3 | 849 | 385.10 |

Table- Optical table masses

|  | Side | Keel | Top | Total |
| --- | --- | --- | --- | --- |
| Weigh (kg) | 98.16 | 90.08 | 385.10 | 573.34 |

Table - Masses distribution (computed using T1100261)

Acceptance Criteria

The Mass budget must be

- 579.1 Kg (cf E1100427)+/-25Kg (5%)

Test result:	Passed:   X   	Failed:         .

## Step 7: Shim thickness

| Lockers | Shim thickness (mil) |
| --- | --- |
| A | 130 |
| B | 123 |
| C | 122 |
| D | 120 |

Table – Shims Thickness

Issues/difficulties/comments regarding this test:

Those shims were changed in order to correct initial tilt found on the optical table.

Test result:	Passed:   X   	Failed:         .

## Step 8: Lockers adjustment

Not recorded but within requirements-cf CPS results later.

Acceptance Criteria

- Vertical and horizontal displacement near the lockers must be lower than 2 mils (0.002”)

Test result:	Passed:   X    	Failed:         .

## Step 9 - Electronics Inventory

| Hardware | LIGO reference | S/N |
| --- | --- | --- |
| Coil driver | D0902744 | S1103308 |
|  |  | S1103332 |
| Anti Image filter | D070081 | S1104572 |
| Anti aliasing filter | D1000269 | S1104621 |
|  |  | S1104638 |
| Interface chassis | D1000067 | S1107434 |
|  |  | S1104638 |

Table - Inventory electronics

Acceptance Criteria

- Inventory is complete

Test result:	Passed:  X    	Failed:         .

## Step 10 – Cables inventory – E1100822

We noticed a few mistakes in our early installation of cables, so here is a corrected inventory

| Cable Connects | Cable S/N |
| --- | --- |
| Part Name | Configuration | Corner 1 | Corner 2 | Corner 3 |
| GS13 | Horizontal | S1104670-S114694 | S1104663-S1104606 | S1104667-S1104592 |
| GS13 | Vertical |  |  |  |
| L4C | Horizontal | N/A | N/A | N/A |
| L4C | Vertical | N/A | N/A | N/A |
| Actuator | Horizontal | S1106907 | S1104751 | S1104499 |
|  | Vertical | S1106674 | S1104606 |  |

Table – Cables inventory

Acceptance Criteria

- Cable inventory completed

- E110082 spreadsheet updated

Test result:	Passed:   X   	Failed:         .

## Step 10 - Set up sensors gap

|  | No mass |
| --- | --- |
| Table locked | ADE boxes on |
| Sensors | Offset (Mean) | Std deviation |
| H1 | 113.73 | 9.7958 |
| H2 | 38.192 | 12.716 |
| H3 | -151.86 | 6.5609 |
| V1 | 418.12 | 15.454 |
| V2 | -638.13 | 35.608 |
| V3 | 318.67 | 7.7891 |

Table – Capacitive position sensor readout after gap set-up

Acceptance criteria:

- All mean values must be lower than 400 cts (a bit less than .0005”).

- All standard deviations below 5 counts.

- No cross talk

Test result:	Passed: X 	Failed:       .

## Step 11 - Check Sensor gaps after the platform release

|  | Table locked | Table unlocked |
| --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Difference |
| H1 | 113.73 | 9.7958 | -611.01 | 724.74 |
| H2 | 38.192 | 12.716 | -85.507 | 123.699 |
| H3 | -151.86 | 6.5609 | -178.95 | 27.09 |
| V1 | 418.12 | 15.454 | -723.53 | 1141.65 |
| V2 | -638.13 | 35.608 | -1311.8 | 673.67 |
| `V3 | 318.67 | 7.7891 | -416.11 | 734.78 |

Table – Sensor gaps after platform release

Acceptance criteria:

- Absolute values of the difference between the unlocked and the locked table must be below:

  - 1600 cts for horizontal sensors (~0.002”)

  - 1600 cts for vertical sensors (~0.002”)

- Considering the acceptance criteria of step 4, all mean values must be lower than

  - 2000 cts for horizontal sensors (~0.0025”)

  - 2000 cts for vertical sensors (~0.0025”)

Test result:	Passed:  X    	Failed:         .

## Step 12 - Position Sensors unlocked/locked Power Spectrum

Data files in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_2/Powerspectra/Undamped

LLO_HAM_ISI_Unit_2_Calibrated_PSD_CPS_GS13_Unlocked_Locked_2011_06_09.mat

Scripts files for processing and plotting in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_2/Powerspectra/Undamped

- Powerspectra_Measurement_HAM_ISI_Locked_Unlocked.m

Figures in SVN at:

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_2/Figures/Powerspectra/Undamped

LLO_HAM_Unit_2_Calibrated_PSD_CPS_Unlocked_Locked_2011_06_09.pdf

CPS calibration:

The CPS power spectrums are calibrated by using a sensitivity of 30.2 nm/count.

Figure - Calibrated CPS power spectrum

Figure – Power spectrum Calibrated GS13

Data files in SVN at: seismic/HAM-ISI/L1/HAM2/Data/Spectra/Undamped/

- LLO_ISI_HAM2_ASD_m_CPS_T240_L4C_GS13_Locked_vs_Unlocked_2012_03_02.mat

Issues/difficulties/comments regarding this test:

H3 CPS seems to have a lower response between 10 and 13 Hz (locked config), this was the only time this was observed, and response appears normal in other configuration.

Scripts files for taking and processing the data, and plotting it in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Plot_ASD_Unlocked_Locked_HAM_ISI.m

- Plot_ASD_Unlocked_Locked_Group_HAM_ISI.m

Figures in SVN at: seismic/HAM-ISI/L1/HAM2/Data/Figures/Spectra/Undamped/

- LLO_ISI_HAM2_ASD_m_GS13_Requirements_Locked_vs_Unlocked_2012_03_02.pdf

- LLO_ISI_HAM2_ASD_m_CPS_Requirements_Locked_vs_Unlocked_2012_03_02.pdf

Acceptance criteria:

- No cross talk (peaks at low frequencies + harmonics on measurements)

- Magnitudes of power spectra must be between requirement curves such as in the following figures (dashed lines)

| Sensors | ISI state | Frequency (Hz) | 2x10-2 | 1x10-1 | 1 | 10 | 20 | 100 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| GS-13 | Table locked | Max | 3x10-1 | 3x10-4 | 3x10-7 | 10-7 |  | 10-11 | 10-14 |
|  |  | Min | 3x10-4 | 3x10-7 | 3x10-10 | 10-12 |  | 10-14 | 10-17 |
|  | Table unlocked | Max | 1 | 3x10-3 | 10-5 | 10-9 |  | 10-11 | 10-14 |
|  |  | Min | 10-4 | 3x10-7 | 10-9 | 10-13 |  | 10-15 | 10-18 |
| CPS | Table locked | Max | 2*10-7 | 2x10-8 | 10-8 | 5x10-8 | 2x10-7 | 5x10-9 | 10-9 |
|  |  | Min | 5x10-9 | 2x10-9 | 8x10-10 | 5x10-10 |  | 10-10 | 5x10-11 |
|  | Table unlocked | Max | 2x10-6 | 8x10-7 | 8x10-7 | 5x10-8 | 2x10-7 | 2x10-8 | 10-9 |
|  |  | Min | 10-7 | 5x10-8 | 8x10-9 | 5x10-10 |  | 2x10-10 | 10-10 |

Table - Step 6 -Normal conditions-Sensors powerspectra requirements

Test result:	Passed:   X   	Failed:          .

## Step 13- GS13 power spectrum -tabled tilted

The figure below presents the GS13 power spectrum when the table is unlocked and loaded with a 20Kg mass at one of its corner.

Figure – Power spectrum Calibrated GS13 with mass at corner

Data files in SVN at: seismic/HAM-ISI/L1/HAM2/Data/Spectra/Undamped/

- LLO_ISI_HAM2_ASD_m_GS13_Stage_Tilted_2012_02_27.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Plot_ASD_Tilted_Stage_HAM_ISI.m

Figures in SVN at: seismic/HAM-ISI/L1/HAM2/Data/Figures/Spectra/Undamped/

- LLO_ISI_HAM2LLO_ISI_HAM2_ASD_m_GS13_Stage_Tilted_2012_02_27_2012_02_27.pdf

Issues/difficulties/comments regarding this test:

One could notice that H3 and V3 differ at high frequency from the other instruments. This is due to the fact that we had wrongly plugged those instruments onto the L4C side of a BSC interface. Therefore those instruments did not benefit from the dewhitening filters as the other instruments did.

Acceptance criteria:

Test result:	Passed:   X   	Failed:          .

## Step 14- GS13 pressure readout

GS-13 pressure sensors constantly read  77 or 78 KPa.

Test result:	Passed:   X   	Failed:          .

## Step 15- Actuators Sign and range of motion (Local drive)

|  | Negative drive | Positive drive |
| --- | --- | --- |
| H1 readout (count) | -23083.026 | 24046.844 |
| H2 readout (count) | -22944.266 | 23368.198 |
| H3 readout (count) | -24318.274 | 24244.29 |
| V1 readout (count) | -19824.716 | 18694.398 |
| V2 readout (count) | -24564.902 | 24903.306 |
| V3 readout (count) | -21513.128 | 21275.082 |

Table - Range of motion - Local drive

Data files in SVN at:  seismic/HAM-ISI/L1/HAM2/Data/Static_Tests/

- LLO_HAM-ISI_Unit_2_Range_Of_Motion_20110609.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Range_Motion_HAM_ISI.m

Acceptance criteria:

- Main couplings sensors readout must be at least 16000 counts (~0.02”)

- A positive offset drive on one actuator must give positive sensor readout on the collocated sensor. Signs will also be tested when measuring local to local transfer functions.

Test result:	Passed:   X   	Failed:         .

## Step 16 - Static Testing (Tests in the local basis)

|  |  | H1 | H2 | H3 | V1 | V2 | V3 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Actuators (1000 counts) | H1 | 1768.52176 | 1120.3965 | 1126.4326 | 2.3293584 | -7.4221 | 33.4536 |
|  | H2 | 1121.76976 | 1822.10976 | 1138.1846 | 10.5282164 | -10.28954 | 9.0632 |
|  | H3 | 1131.02036 | 1154.148 | 1849.87592 | 18.5308684 | -32.36534 | 36.4462 |
|  | V1 | 144.12706 | 161.86032 | -309.1136 | 1218.91386 | -35.6438 | -507.9544 |
|  | V2 | -347.93947 | 172.49 | 167.3648 | -534.30694 | 1243.82962 | -12.4158 |
|  | V3 | 144.38038 | -341.69618 | 173 | -29.688582 | -582.03506 | 1348.44654 |

Table - Main and cross coupling

Data files in SVN at:  seismic/HAM-ISI/L1/HAM2/Data/Static_Tests/

- / LLO_HAM_ISI_Unit_2_Sensor_Readout_Local_20110609.mat

Scripts files for taking data in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Static_Test_Local_Basis_HAM_ISI.m

Issues/difficulties/comments regarding this test:

V1 and V2 appear to give responses just lower from the criteria.

Acceptance criteria:

- Vertical

For a +1000 count offset drive on vertical actuators

  - Collocated sensors must be 1400 counts +/- 10%

- Horizontal

For a +1000 count offset drive on horizontal actuators

  - Collocated sensors must be 2000 counts +/- 10%

  - Non-collocated horizontal sensors must be 1250 counts +/-10%

Test result:	Passed:       	Failed:    X     .

## Step 17 - Linearity test

|  | Slope | Average slope | Variation from average(%) |
| --- | --- | --- | --- |
| H1 | 1.84147738 | 1.8424 | -0.09879621 |
| H2 | 1.83885246 |  | -0.36128864 |
| H3 | 1.84706619 |  | 0.46008485 |
| V1 | 1.33770934 | 1.3335 | 0.41841887 |
| V2 | 1.32772348 |  | -0.58016727 |
| V3 | 1.33514264 |  | 0.1617484 |

Table - Slopes and offset of the triplet Actuators - HAM-ISI - Sensors

Scripts files for taking data in SVN at: seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Linearity_Test_Awgstream_HAM_ISI.m

Figure - Horizontal and vertical actuators x HAM-ISI x sensors

Data files in SVN at: seismic/HAM-ISI/L1/HAM2/Data/Linearity_Test/

- LLO_HAM_ISI_Unit_2_Linearity_test_20110609.mat

Figures in SVN at: seismic/HAM-ISI/L1/HAM2/Data/Figures/Linearity_Test/

- LLO_HAM_ISI_Unit_2_Linearity_test_20110609.fig

LLO_HAM_ISI_Unit_2_Linearity_test_20110609.pdf

Acceptance criteria:

- Horizontal and vertical slopes of the triplet actuators x HAM-ISI x sensors:  Average slope +/- 1%

Test result:	Passed:   X   	Failed:         .

## Step 18 - Cartesian Basis Static Testing

|  | 1000 counts drive | X Drive | Y Drive | Z Drive | Rx Drive | Ry Drive | Rz Drive |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Sensors readout (count) | H1 | 263.528 | -390.4432 | 39.232 | -351.599 | -234.314 | -1870.593 |
|  | H2 | 232.73 | 510.05 | 51.46 | 511.84 | -214.09 | -1926.44 |
|  | H3 | -492.32 | 23.53 | 10.56 | 70.06 | 532.44 | -1901.82 |
|  | V1 | -5.871 | 6.292 | 248.899 | -510.236 | -1619.426 | 11.019 |
|  | V2 | -21.28 | -33.566 | 239.421 | 1633.514 | 398.43 | -57.855 |
|  | V3 | 2.8 | -18.2 | 270.36 | -1169.8 | 1208.911 | 29.8 |
|  | Direction read out | 492.38 | 524.71 | 256.965 | 2516.66 | 2506.73 | 2404.763 |

Table - Tests in the general coordinate basis

Issues/difficulties/comments regarding this test:

This is the first ISI oriented in the Y direction. Matrices were redesigned for that purpose.

Acceptance criteria:

For a positive drive in the Cartesian basis:

- Local sensor readout must have the same sign that the reference table (CONT2ACT check)

- Cartesian sensors read out must be positive (DISP2CEN check) in the drive direction

Test result:	Passed:      	Failed:         .

## Step 19- Frequency response

## Step 19.1 - Local to local measurements

Local to local transfer functions have been measured with 90 repetitions.

Data files in SVN at:

Local to local transfer functions have been measured with 90 repetitions.

Data files in SVN at:

seismic/HAM-ISI/L1/HAM2/Data/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_HAM3_Data_TF_L2L_50mHz_500mHz_20120201-000859.mat

- LLO_ISI_HAM3_Data_TF_L2L_500mHz_5Hz_20120131-213926.mat

- LLO_ISI_HAM3_Data_TF_L2L_5Hz_200Hz_20120131-205053.mat

- LLO_ISI_HAM3_Data_TF_L2L_200Hz_800Hz_20120131-191720.mat

Data collection script files: seismic/HAM-ISI/L1/HAM2/Scripts/Data_Collection/

- Run_Exc_Batch_L1ISIHAM2.m

Scripts files for processing and plotting in SVN at:

seismic/HAM-ISI/L1/HAM2/Scripts/Control_Scripts/

- Step_1_TF_L2L_L1_ISI_HAM3.m

Figures in SVN at:

seismic/HAM-ISI/L1/HAM2/Data/Figures/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_HAM3_TF_L2L_Raw_from_ACT_to_CPS_2012_01_31.fig

- LLO_ISI_HAM3_TF_L2L_Raw_from_ACT_to_GS13_2012_01_31.fig

Storage of measured transfer functions in the SVN at:

seismic/HAM-ISI/L1/HAM2/Data/Transfer_Functions/Measurements/Undamped/

- LLO_HAM_ISI_Unit_3_Data_TF_L2L_2011_04_06.mat

The local to local transfer functions are presented below.

Figure - Local to Local Measurements – Capacitive sensors

Figure - Local to Local Measurements – Inertial sensors

Zoom around 1 Hz of the Inertial sensors response.

Issues/difficulties/comments regarding this test:

We can see that H1 has a different response from H2, and H3, whereas V2 differs slightly from V1 and V3.

If we compare to testing right after assembly, we can see somewhat similar features except H3 is different from the others. Now, all sensors were changed between the assembly and the side-chamber testing. It so happened that the one which used to be in H3 is now in H1, but all other sensors are different.

Figure - Local to Local Measurements – Horizontal Inertial sensors

Figure - Local to Local Measurements – Vertical Inertial sensors

Figure - Local to Local Measurements – Inertial sensors comparison with HAM 3 (dashed)

Finally, when we extract the horizontal seismometers responses and compare it with their huddle tests, we find that we can see the most of the difference between sensors.

Acceptance criteria:

- Local to local measurements

  - On CPS, the phase must be 0º at DC

  - On Geophones, the phase must be -90º at DC

  - Identical shape in each corner

- Cartesian to Cartesian measurements

  - On CPS, the phase must be 0º at DC

  - On Geophones, the phase must be -90º at DC

  - Identical shape X/Y and RX/RY

Test result:	Passed:   X   	Failed:         .

## Step 20 - Lower Zero Moment Plane

Since we did not pass one requirement (Y offset, now become X offset) during initial testing, we did re-test it.

X offset 2.465 mm

Y offset 0.947 mm

Issues/difficulties/comments regarding this test:

We can notice that we have a fairly low coherence on both measurements and that additional testing may be useful later.

Acceptance criteria:

- X offset must be less than 2 mm

- Y offset must be less than 2 mm

Test result:	Passed:       	Failed:   X      .

Conclusion Side Chamber testing

Only a few issues were found during the testing of this unit, which are summed up here:

- Local 2 local transfer functions displays a slightly different behavior around 1 Hz in corner 3 than in the others. After further study this appears to be due to the sensors themselves.

- X LZMP measurements do not meet requirements, we can conduct more tests once in chamber.

- V1 and V2 local to local static testing barely missed requirements.

H3 CPS displayed a low answer in one test, this is something we can investigate later.

# INITIAL CHAMBER TESTING

The HAM-ISI was installed in HAM 2 on March 5th 2012, loaded on March 6th and connected to the electronics on March 8th. We initially had to adjust its level using HEPI (indeed dropping the ISI in there lowered the support tubes by about 40 mils). Additionally the locker shims were slightly adjusted again to correct the level. CPS sensors were zeroed again.

Testing mostly happened between March 9th and March 15th.

## Step 5: Check level of Stage 1 Optical Table

Acceptance Criteria

- The maximum angle of the table with the horizontal mustn’t exceed  ~100µrad

This table appears to have local deformation but no excessive tilt

Test result:	Passed:   X  	Failed:         .

## Step 3: Blade spring profile

Re-measured because we had to change shims to meet the platform level requirements.

| Blade # | Base (") | Tip(") | Flatness (mils) |
| --- | --- | --- | --- |
| 1 | .4829 | .4712 | +11.7 |
| 2 | .4895 | .4730 | +16.5 |
| 3 | .4834 | .4691 | +14.3 |

Table 1 - Blade profile

Issues/difficulties/comments regarding this test:

Those were re-measured because we adjusted the shims to get the table within level requirements.

Acceptance Criteria:

- Blades must be flat within 0.015" inches.

Test result:	Passed:      	Failed:  X       .

## Step 4: Gap checks on actuators-after installation on Stage 1

Not recorded but checked and within requirements.

Acceptance Criteria

- Gaps must be within 0.010” of design value (0.080”)

Test result:	Passed:   X  	Failed:         .

## Step 6: Mass budget

Figure – Keel Masses and Wall masses location

|  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| k1 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k2 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
| k3 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k4 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
| k5 |  |  |  |  | 1 |  | 1 | 35.1 | 15.92 |
| k6 | 1 | 1 | 1 |  |  |  | 1 | 31.1 | 14.11 |
|  | 3 | 3 | 3 | 0 | 3 | 0 | 6 | 198.6 | 90.08 |

Table- Keel masses

|  | 00 | 01 | 02 | 03 | 04 | 05 | 06 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0.6 | 1.1 | 2.2 | 4.5 | 7.9 | 15.6 | 27.2 | lbs | kgs |
| w9 | 1 | 1 | 1 | 1 |  | 1 |  | 24 | 10.89 |
| w1 | 1 | 1 | 2 | 1 |  | 1 |  | 26.2 | 11.88 |
| w2 | 3 |  | 4 |  |  | 1 |  | 26.2 | 11.88 |
| w3 | 2 |  | 1 | 1 |  | 1 |  | 23.5 | 10.66 |
| w4 | 2 | 1 |  | 1 |  | 1 |  | 22.4 | 10.16 |
| w5 |  |  | 1 | 1 |  | 1 |  | 22.3 | 10.12 |
| w6 | 2 | 1 | 1 | 1 |  | 1 |  | 24.6 | 11.16 |
| w7 | 1 |  |  | 1 |  | 1 |  | 20.7 | 9.39 |
| w8 | 2 |  | 2 | 1 |  | 1 |  | 25.7 | 11.66 |
| Side Masses Total | 14 | 4 | 12 | 8 | 0 | 9 | 0 | 215.6 | 97.79 |

Table- Side masses

|  | D972213 | D972214 | D972215 | D0901075 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | 2.5 kg | 5 kg | 10 kg |  |  |
|  | 610 | 375 | 230 | 5.5 | 11 | 22 | lbs | kgs |
| Top Masses | 0 | 2 | 0 | 0 | 3 | 3 | 849 | 385.10 |

Table- Optical table masses

|  | Side | Keel | Top | Total |
| --- | --- | --- | --- | --- |
| Weigh (kg) | 97.79 | 90.08 | 385.10 | 572.98 |

Table - Masses distribution (computed using T1100261)

Acceptance Criteria

The Mass budget must be

- 579.1 Kg (cf E1100427)+/-25Kg (5%)

Test result:	Passed:   X   	Failed:         .

## Step 7: Shim thickness

| Lockers | Shim thickness (mil) |
| --- | --- |
| A | 130 |
| B | 120 |
| C | 128 |
| D | 120 |

Table – Shims Thickness

Issues/difficulties/comments regarding this test:

Those shims were changed in order to correct initial tilt found on the optical table.

Test result:	Passed:   X   	Failed:         .

## Step 8: Lockers adjustment

Not recorded but within requirements-cf CPS results later.

Acceptance Criteria

- Vertical and horizontal displacement near the lockers must be lower than 2 mils (0.002”)

Test result:	Passed:   X    	Failed:         .

## Step 9 - Electronics Inventory

| Hardware | LIGO reference | S/N |
| --- | --- | --- |
| Coil driver | D0902744 | S1103308 |
|  |  | S1103332 |
| Anti Image filter | D070081 | S1104572 |
| Anti aliasing filter | D1000269 | S1104621 |
|  |  | S1104638 |
| Interface chassis | D1000067 | S1107434 |
|  |  | S1107435 |

Table - Inventory electronics

Note that we replaced the BSC-ISI interface that we used for side-chamber testing with the HAM-ISI interface that had gotten replaced.

Acceptance Criteria

- Inventory is complete

Test result:	Passed:  X    	Failed:         .

## Step 10 - Set up sensors gap

|  | No mass |
| --- | --- |
| Table locked | ADE boxes on |
| Sensors | Offset (Mean) | Std deviation |
| H1 | 248.84 | 6.5019 |
| H2 | -262.73 | 5.0111 |
| H3 | 193.05 | 6.3325 |
| V1 | -106.2 | 6.7945 |
| V2 | -223.67 | 9.3371 |
| V3 | -426.2 | 5.3248 |

Table – Capacitive position sensor readout after gap set-up

Acceptance criteria:

- All mean values must be lower than 400 cts (a bit less than .0005”).

- All standard deviations below 5 counts.

- No cross talk

Test result:	Passed: X 	Failed:       .

## Step 11 - Check Sensor gaps after the platform release

|  | Table locked | Table unlocked |
| --- | --- | --- |
| Sensors | Offset (Mean) | Std deviation | Offset (Mean) | Difference |
| H1 | 248.84 | 6.5019 | 1165.1 | -916.26 |
| H2 | -262.73 | 5.0111 | -286.86 | 24.13 |
| H3 | 193.05 | 6.3325 | 733.42 | -540.37 |
| V1 | -106.2 | 6.7945 | -1433.3 | 1327.1 |
| V2 | -223.67 | 9.3371 | 885.55 | -1109.22 |
| `V3 | -426.2 | 5.3248 | -3018.1 | 2591.9 |

Table – Sensor gaps after platform release

Acceptance criteria:

- Absolute values of the difference between the unlocked and the locked table must be below:

  - 1600 cts for horizontal sensors (~0.002”)

  - 1600 cts for vertical sensors (~0.002”)

- Considering the acceptance criteria of step 4, all mean values must be lower than

  - 2000 cts for horizontal sensors (~0.0025”)

  - 2000 cts for vertical sensors (~0.0025”)

Test result:	Passed:  X    	Failed:         .

## Step 12 - Position Sensors unlocked/locked Power Spectrum

Data files in SVN at: seismic/HAM-ISI/L1/HAM2/Data/Spectra/Undamped

- LLO_HAM_ISI_Unit_2_Calibrated_PSD_CPS_GS13_Unlocked_Locked_2012_03_14.mat

Figures in SVN at: seismic/HAM-ISI/L1/HAM2/Data/Figures/Spectra/Undamped

/opt/svncommon/seisvn/seismic/HAM-ISI/X2/Data/Unit_2/Figures/Powerspectra/Undamped

- LLO_HAM_Unit_2_Calibrated_PSD_CPS_Unlocked_Locked_2012_03_14.pdf

CPS calibration:

The CPS power spectrums are calibrated by using a sensitivity of 30.2 nm/count.

Figure - Calibrated CPS power spectrum

Figure – Power spectrum Calibrated GS13

Data files in SVN at: seismic/HAM-ISI/L1/HAM2/Data/Spectra/Undamped/

- LLO_ISI_HAM2_ASD_m_CPS_T240_L4C_GS13_Locked_vs_Unlocked_2012_03_14.mat

Issues/difficulties/comments regarding this test:

H3 CPS seems to have a lower response between 10 and 13 Hz (locked config), this was the only time this was observed, and response appears normal in other configuration.

Scripts files for taking and processing the data, and plotting it in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Plot_ASD_Unlocked_Locked_HAM_ISI.m

- Plot_ASD_Unlocked_Locked_Group_HAM_ISI.m

Figures in SVN at: seismic/HAM-ISI/L1/HAM2/Data/Figures/Spectra/Undamped/

- LLO_ISI_HAM2_ASD_m_GS13_Requirements_Locked_vs_Unlocked_2012_03_14.pdf

- LLO_ISI_HAM2_ASD_m_CPS_Requirements_Locked_vs_Unlocked_2012_03_14.pdf

Acceptance criteria:

- No cross talk (peaks at low frequencies + harmonics on measurements)

- Magnitudes of power spectra must be between requirement curves such as in the following figures (dashed lines)

| Sensors | ISI state | Frequency (Hz) | 2x10-2 | 1x10-1 | 1 | 10 | 20 | 100 | 1000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| GS-13 | Table locked | Max | 3x10-1 | 3x10-4 | 3x10-7 | 10-7 |  | 10-11 | 10-14 |
|  |  | Min | 3x10-4 | 3x10-7 | 3x10-10 | 10-12 |  | 10-14 | 10-17 |
|  | Table unlocked | Max | 1 | 3x10-3 | 10-5 | 10-9 |  | 10-11 | 10-14 |
|  |  | Min | 10-4 | 3x10-7 | 10-9 | 10-13 |  | 10-15 | 10-18 |
| CPS | Table locked | Max | 2*10-7 | 2x10-8 | 10-8 | 5x10-8 | 2x10-7 | 5x10-9 | 10-9 |
|  |  | Min | 5x10-9 | 2x10-9 | 8x10-10 | 5x10-10 |  | 10-10 | 5x10-11 |
|  | Table unlocked | Max | 2x10-6 | 8x10-7 | 8x10-7 | 5x10-8 | 2x10-7 | 2x10-8 | 10-9 |
|  |  | Min | 10-7 | 5x10-8 | 8x10-9 | 5x10-10 |  | 2x10-10 | 10-10 |

Table - Step 6 -Normal conditions-Sensors powerspectra requirements

Test result:	Passed:   X   	Failed:          .

## Step 13- GS13 power spectrum -tabled tilted

The figure below presents the GS13 power spectrum when the table is unlocked and loaded with a 20 Kg mass at one of its corner.

Figure – Power spectrum Calibrated GS13 with mass at corner

Data files in SVN at: seismic/HAM-ISI/L1/HAM2/Data/Spectra/Undamped/

- LLO_ISI_HAM2_ASD_m_GS13_Stage_Tilted_2012_03_12.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Plot_ASD_Tilted_Stage_HAM_ISI.m

Figures in SVN at: seismic/HAM-ISI/L1/HAM2/Data/Figures/Spectra/Undamped/

- LLO_ISI_HAM2LLO_ISI_HAM2_ASD_m_GS13_Stage_Tilted_2012_03_12.pdf

Acceptance criteria:

Test result:	Passed:   X   	Failed:          .

## Step 14- GS13 pressure readout

GS-13 pressure sensors constantly read  77 or 78 KPa.

Test result:	Passed:   X   	Failed:          .

## Step 15- Actuators Sign and range of motion (Local drive)

|  | Negative drive | Positive drive |
| --- | --- | --- |
| H1 readout (count) | -23073.888 | 24341.642 |
| H2 readout (count) | -23112.536 | 23447.74 |
| H3 readout (count) | -24741.966 | 24378.672 |
| V1 readout (count) | -19926.948 | 19208.934 |
| V2 readout (count) | -24683.802 | 25557.282 |
| V3 readout (count) | -22756.788 | 21081.596 |

Table - Range of motion - Local drive

Data files in SVN at:  seismic/HAM-ISI/L1/HAM2/Data/Static_Tests/

- LLO_HAM-ISI_Unit_2_Range_Of_Motion_20120312.mat

Scripts files for taking and processing the data, and plotting it in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Range_Motion_HAM_ISI.m

Acceptance criteria:

- Main couplings sensors readout must be at least 16000 counts (~0.02”)

- A positive offset drive on one actuator must give positive sensor readout on the collocated sensor. Signs will also be tested when measuring local to local transfer functions.

Test result:	Passed:   X   	Failed:         .

## Step 16 - Static Testing (Tests in the local basis)

|  |  | H1 | H2 | H3 | V1 | V2 | V3 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Actuators (1000 counts) | H1 | 1815.7554 | 1140.44682 | 1147.0721 | -13.442 | -12.99658 | -1.7162 |
|  | H2 | 1135.0146 | 1847.06036 | 1165.68858 | -12.4966 | 9.92384 | -2.9238 |
|  | H3 | 1144.9278 | 1179.62188 | 1874.04406 | -1.9248 | -1.1169 | 1.4338 |
|  | V1 | 156.3542 | 181.078276 | -317.6344 | 1300.34141 | -35.5098 | -577.112 |
|  | V2 | -356.0164 | 182.38216 | 171.691251 | -585.6892 | 1327.21362 | -29.607 |
|  | V3 | 161.9512 | -361.92448 | 185.35227 | -41.395 | -615.21274 | 1402.104 |

Table - Main and cross coupling

Data files in SVN at:  seismic/HAM-ISI/L1/HAM2/Data/Static_Tests/

- / LLO_HAM_ISI_Unit_2_Sensor_Readout_Local_20120312.mat

Scripts files for taking data in SVN at:

seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Static_Test_Local_Basis_HAM_ISI.m

Issues/difficulties/comments regarding this test:

V1 and V2 appear to again give responses just lower from the criteria.

Acceptance criteria:

- Vertical

For a +1000 count offset drive on vertical actuators

  - Collocated sensors must be 1400 counts +/- 10%

- Horizontal

For a +1000 count offset drive on horizontal actuators

  - Collocated sensors must be 2000 counts +/- 10%

  - Non-collocated horizontal sensors must be 1250 counts +/-10%

Test result:	Passed:       	Failed:    X     .

## Step 17 - Linearity test

|  | Slope | Average slope | Variation from average(%) |
| --- | --- | --- | --- |
| H1 | 1.8263 | 1.8454 | -1.03500596 |
| H2 | 1.8424 |  | -0.16256638 |
| H3 | 1.8675 |  | 1.19757234 |
| V1 | 1.3215 | 1.3497 | -2.09177121 |
| V2 | 1.3255 |  | -1.79541638 |
| V3 | 1.4022 |  | 3.88718759 |

Table - Slopes and offset of the triplet Actuators - HAM-ISI - Sensors

Scripts files for taking data in SVN at: seismic/HAM-ISI/Common/Testing_Functions_HAM_ISI/

- Linearity_Test_Awgstream_HAM_ISI.m

Figure - Horizontal and vertical actuators x HAM-ISI x sensors

Data files in SVN at: seismic/HAM-ISI/L1/HAM2/Data/Linearity_Test/

- LLO_HAM_ISI_Unit_2_Linearity_test_20120315.mat

Figures in SVN at: seismic/HAM-ISI/L1/HAM2/Data/Figures/Linearity_Test/

- LLO_HAM_ISI_Unit_2_Linearity_test_20120315.fig

LLO_HAM_ISI_Unit_2_Linearity_test_20120315.pdf

Acceptance criteria:

- Horizontal and vertical slopes of the triplet actuators x HAM-ISI x sensors:  Average slope +/- 1%

Test result:	Passed:      	Failed:   X      .

## Step 19- Frequency response

## Step 19.1 - Local to local measurements

Local to local transfer functions have been measured with 90 repetitions.

Data files in SVN at:

Local to local transfer functions have been measured with 90 repetitions.

Data files in SVN at:

seismic/HAM-ISI/L1/HAM2/Data/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_HAM3_Data_TF_L2L_50mHz_500mHz_20120201-000859.mat

- LLO_ISI_HAM3_Data_TF_L2L_500mHz_5Hz_20120131-213926.mat

- LLO_ISI_HAM3_Data_TF_L2L_5Hz_200Hz_20120131-205053.mat

- LLO_ISI_HAM3_Data_TF_L2L_200Hz_800Hz_20120131-191720.mat

Data collection script files: seismic/HAM-ISI/L1/HAM2/Scripts/Data_Collection/

- Run_Exc_Batch_L1ISIHAM2.m

Scripts files for processing and plotting in SVN at:

seismic/HAM-ISI/L1/HAM2/Scripts/Control_Scripts/

- Step_1_TF_L2L_L1_ISI_HAM3.m

Figures in SVN at:

seismic/HAM-ISI/L1/HAM2/Data/Figures/Transfer_Functions/Measurements/Undamped/

- LLO_ISI_HAM3_TF_L2L_Raw_from_ACT_to_CPS_2012_01_31.fig

- LLO_ISI_HAM3_TF_L2L_Raw_from_ACT_to_GS13_2012_01_31.fig

Storage of measured transfer functions in the SVN at:

seismic/HAM-ISI/L1/HAM2/Data/Transfer_Functions/Measurements/Undamped/

- LLO_HAM_ISI_Unit_3_Data_TF_L2L_2011_04_06.mat

The local to local transfer functions are presented below.

TBC

Figure - Local to Local Measurements – Capacitive sensors

Figure - Local to Local Measurements – Inertial sensors

Acceptance criteria:

- Local to local measurements

  - On CPS, the phase must be 0º at DC

  - On Geophones, the phase must be -90º at DC

  - Identical shape in each corner

- Cartesian to Cartesian measurements

  - On CPS, the phase must be 0º at DC

  - On Geophones, the phase must be -90º at DC

  - Identical shape X/Y and RX/RY

Test result:	Passed:   X   	Failed:         .

Conclusion Initial In-Chamber testing

Most minor issues found on this ISI during side-chamber testing were confirmed during in chamber tests.

- Local 2 local transfer functions displays a slightly different behavior around 1 Hz in corner 3 than in the others. We still believe this to be due to the sensors themselves.

- X LZMP measurements do not meet requirements, we can conduct more tests once in chamber.

- V1 and V2 local to local static testing barely missed requirements.

- H3 CPS displayed a low answer in one test, this is something we can investigate later.

Surprisingly, the linearity test, which had given good results during side-chamber testing fails here.