*LIGO Laboratory / LIGO Scientific Collaboration*

LIGO-E1200226-v6 Advanced LIGO 6/26/2017

TwinCAT Library for Low Noise VCO

Daniel Sigg

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note  
of the LIGO Laboratory.

|  |  |
| --- | --- |
| **California Institute of Technology**  **LIGO Project – MS 18-34**  **1200 E. California Blvd.**  **Pasadena, CA 91125**  Phone (626) 395-2129  Fax (626) 304-9834  E-mail: info@ligo.caltech.edu | **Massachusetts Institute of Technology**  **LIGO Project – NW22-295**  **185 Albany St**  **Cambridge, MA 02139**  Phone (617) 253-4824  Fax (617) 253-7014  E-mail: info@ligo.mit.edu |
| **LIGO Hanford Observatory**  **P.O. Box 159**  **Richland WA 99352**  Phone 509-372-8106  Fax 509-372-8137 | **LIGO Livingston Observatory**  **P.O. Box 940**  **Livingston, LA 70754**  Phone 225-686-3100  Fax 225-686-7189 |

http://www.ligo.caltech.edu/

|  |  |
| --- | --- |
| **Library** | |
| Title | LowNoiseVco |
| Version | 6 |
| TwinCAT version | 2.11 |
| Name space | – |
| Author | Daniel Sigg |
| Description | Controls the low noise VCO, [D0900605](https://dcc.ligo.org/cgi-bin/private/DocDB/ShowDocument?.submit=Number&docid=D0900605&version=), and the frequency difference mixer, [D1600499](https://dcc.ligo.org/LIGO-D1600499).  The low noise VCO is based on a frequency difference divider. It requires a 71MHz/10dBm reference source and a VCO source at either 125MHz or 79MHz. Both RF levels as well as the RF level at the output of the frequency difference divider are monitored. The only set value is an offset into the VCO which translates into a frequency offset at the output. A binary output is used to enable the excitation input. Additional monitors are available for the tune voltage, the state of the excitation switch, and a power ok bit.  If a frequency counter has been setup through the timing system, the measured frequency can be stabilized by feeding back to the bias offset. This then allows the user to select a fixed output frequency.  The frequency difference mixer is using the same RF mixer circuit but without a divider and a VCO. It implements none of the extra frequency controls of the VCO neither.  The 3 RF power monitors which have the calibration  The corresponding temperature readout has the calibration  The RF power levels should be alarmed when outside ±1dBm of nominal. |
| Error codes | Low Noise VCO:  0x01 – Power supply voltages out-of-range  0x02 – Reference RF power level out-of-range  0x04 – Divider RF power level out-of-range  0x08 – Output RF power level out-of-range  0x10 – Excitation switch enabled  0x20 – Invalid frequency  0x40 – Controls error  Frequency difference mixer:  0x01 – Power supply voltages out-of-range  0x02 – Reference RF power level out-of-range  0x04 – Input RF power level out-of-range  0x08 – Output RF power level out-of-range  Controls errors:  0x01 – Unity gain frequency too high  0x02 – Unity gain frequency too low  0x04 – High limit reached  0x08 – Low limit reached  0x10 – Invalid error signal  0x20 – Invalid set frequency |
| Library dependencies: | Error, SaveRestore, ReadADC. WriteDAC |

|  |  |
| --- | --- |
| **Hardware Input Type**  TYPE LowNoiseVcoInStruct :  STRUCT  ReferenceMon: INT;  DividerMon: INT;  OutputMon: INT;  ReferenceTemp: INT;  DividerTemp: INT;  OutputTemp: INT;  TuneMon: INT;  Frequency: LREAL; (\* not used \*)  ExcitationSwitch: BOOL;  PowerOk: BOOL;  FrequencyLive: BOOL; (\* not used \*)  END\_STRUCT  END\_TYPE | |
| Type name | LowNoiseVcoInStruct |
| Description | Structure of the hardware inputs that are wired up for the low noise VCO |
| Definition | STRUCT |
| Element | Name: ReferenceMon  Type: INT  Description: Monitors the RF power at the reference input |
| Element | Name: DividerMon  Type: INT  Description: Monitors the RF power at the divider input |
| Element | Name: OutputMon  Type: INT  Description: Monitors the RF power after the output amplifier |
| Element | Name: ReferenceTemp  Type: INT  Description: Monitors the temperature of the reference RF detector |
| Element | Name: DividerTemp  Type: INT  Description: Monitors the temperature of the divider RF detector |
| Element | Name: OutputTemp  Type: INT  Description: Monitors the temperature of the output RF detector |
| Element | Name: TuneMon  Type: INT  Description: Monitor for the frequency offset |
| Element | Name: Frequency  Type: LREAL  Description: Measured frequency |
| Element | Name: ExcitationSwitch  Type: BOOL  Description: Monitors the excitation input enable |
| Element | Name: PowerOk  Type: BOOL  Description: Voltage monitor readback |
| Element | Name: FrequencyLive  Type: BOOL  Description: Keep alive for frequency measurement |

|  |  |
| --- | --- |
| **Hardware Output Type**  TYPE LowNoiseVcoOutStruct :  STRUCT  TuneOfs: INT;  ExcitationEn: BOOL;  END\_STRUCT  END\_TYPE | |
| Type name | LowNoiseVcoOutStruct |
| Description | Structure of the hardware outputs that are wired up for the low noise VCO |
| Definition | STRUCT |
| Element | Name: TuneOfs  Type: INT  Description: Setpoint for the frequency offset |
| Element | Name: ExcitationEn  Type: BOOL  Description: Enables the excitation input |

|  |  |
| --- | --- |
| **User Interface Type**  TYPE LowNoiseVcoStruct :  STRUCT  Error: ErrorStruct;  ReferenceMon: LREAL;  ReferenceNom: LREAL;  DividerMon: LREAL;  DividerNom: LREAL;  OutputMon: LREAL;  OuptutNom: LREAL;  ReferenceTemp: LREAL;  DividerTemp: LREAL;  OutputTemp: LREAL;  TuneOfs: LREAL;  TuneMon: LREAL;  TuneLimit: LREAL;  ExcitationSwitch: BOOL;  ExcitationEn: BOOL;  PowerOk: BOOL;  Frequency: LREAL;  FrequencyFault: BOOL;  Controls: LowNoiseVcoControlsStruct;  END\_STRUCT  END\_TYPE | |
| Type name | LowNoiseVcoStruct |
| Description | Structure of the user interface tags that are used to control the low noise VCO |
| Definition | STRUCT |
| Output Tag | Name: Error  Type: ErrorStruct  Description: For error handler |
| Output Tag | Name: ReferenceMon  Type: LREAL  Description: Monitors the RF power at the reference input in dBm |
| Input Tag | Name: ReferenceNom  Type: LREAL  Description: Nominal value for the RF power at the reference input in dBm |
| Output Tag | Name: DividerMon  Type: LREAL  Description: Monitors the RF power at the divider input in dBm |
| Input Tag | Name: DividerNom  Type: LREAL  Description: Nominal value for the RF power at the divider input in dBm |
| Output Tag | Name: OutputMon  Type: LREAL  Description: Monitors the RF power after the output amplifier dBm |
| Input Tag | Name: OutputNom  Type: LREAL  Description: Nominal value for the RF power at the output amplifier in dBm |
| Output Tag | Name: ReferenceTemp  Type: LREAL  Description: Monitors the temperature of the reference RF detector in C |
| Output Tag | Name: DividerTemp  Type: LREAL  Description: Monitors the temperature of the divider RF detector in C |
| Output Tag | Name: OutputTemp  Type: LREAL  Description: Monitors the temperature of the output RF detector in C |
| Input Tag | Name: TuneOfs  Type: LREAL  Description: Setpoint for the frequency offset in V |
| Output Tag | Name: TuneMon  Type: LREAL  Description: Monitor for the frequency offset in V |
| Input Tag | Name: TuneLimit  Type: LREAL  Description: Limit for the frequency offset in V |
| Input Tag | Name: ExcitationEn  Type: BOOL  Description: Enables the excitation input |
| Output Tag | Name: ExcitationSwitch  Type: BOOL  Description: Monitors the excitation input enable |
| Output Tag | Name: PowerOk  Type: BOOL  Description: Voltage monitor readback |
| Output Tag | Name: Frequency  Type: LREAL  Description: Frequency of the VCO output |
| Output Tag | Name: FrequencyFault  Type: BOOL  Description: Indicates if the frequency of the VCO is no longer updating correctly |
| Input Tag | Name: Controls  Type: LowNoiseVcoControlsStruct  Description: VCO frequency controls parameters |

|  |  |
| --- | --- |
| **User Interface Type**  TYPE LowNoiseVcoControlsStruct:  STRUCT  Error: ErrorStruct;  Fault: BOOL;  SetFrequency: LREAL;  SetFrequencyOffset: LREAL;  DiffFrequency: LREAL;  Enable: BOOL;  UnityGain: LREAL;  ClearInt: BOOL;  END\_STRUCT  END\_TYPE | |
| Type name | LowNoiseVcoControlsStruct |
| Description | Structure of the user interface that is used to control the frequency of the low noise VCO |
| Definition | STRUCT |
| Output Tag | Name: Error  Type: ErrorStruct  Description: For error handler |
| Output Tag | Name: Fault  Type: BOOL  Description: Indicated a servo fault |
| Input Tag | Name: SetFrequency  Type: LREAL  Description: Set frequency in Hz |
| Input Tag | Name: SetFrequencyOffset  Type: LREAL  Description: Set frequency offset in Hz |
| Output Tag | Name: DiffFrequency  Type: LREAL  Description: Difference between measured and set frequency in Hz |
| Input Tag | Name: Enable  Type: BOOL  Description: Enable the servo |
| Input Tag | Name: UnityGain  Type: LREAL  Description: Unity gain frequency in Hz |
| Input Tag | Name: ClearInt  Type: BOOL  Description: Clear the history of the integrator |

|  |  |
| --- | --- |
| **User Interface Type**  TYPE LowNoiseVcoTypeEnum: (VCO, FDD);  END\_TYPE | |
| Type name | LowNoiseVcoTypeEnum |
| Description | Enumerated type to describe the type of the low noise VCO |
| Definition | ENUM |
| Enum Tag | Name: VCO  Description: Standard VCO |
| Enum Tag | Name: FDD  Description: Frequency-difference divider |

|  |  |
| --- | --- |
| **User Interface Type**  TYPE FrequencyDifferenceMixerStruct :  STRUCT  Error: ErrorStruct;  ReferenceMon: LREAL;  ReferenceNom: LREAL;  InputMon: LREAL;  InputNom: LREAL;  OutputMon: LREAL;  OuptutNom: LREAL;  ReferenceTemp: LREAL;  InputTemp: LREAL;  OutputTemp: LREAL;  PowerOk: BOOL;  END\_STRUCT  END\_TYPE | |
| Type name | FrequencyDifferenceMixerStruct |
| Description | Structure of the user interface tags that are used to control the frequency difference mixer |
| Definition | STRUCT |
| Output Tag | Name: Error  Type: ErrorStruct  Description: For error handler |
| Output Tag | Name: ReferenceMon  Type: LREAL  Description: Monitors the RF power at the reference input in dBm |
| Input Tag | Name: ReferenceNom  Type: LREAL  Description: Nominal value for the RF power at the reference input in dBm |
| Output Tag | Name: InputMon  Type: LREAL  Description: Monitors the RF power at the input in dBm |
| Input Tag | Name: InputNom  Type: LREAL  Description: Nominal value for the RF power at the input in dBm |
| Output Tag | Name: OutputMon  Type: LREAL  Description: Monitors the RF power after the output amplifier dBm |
| Input Tag | Name: OutputNom  Type: LREAL  Description: Nominal value for the RF power at the output amplifier in dBm |
| Output Tag | Name: ReferenceTemp  Type: LREAL  Description: Monitors the temperature of the reference RF detector in C |
| Output Tag | Name: InputTemp  Type: LREAL  Description: Monitors the temperature of the input RF detector in C |
| Output Tag | Name: OutputTemp  Type: LREAL  Description: Monitors the temperature of the output RF detector in C |
| Output Tag | Name: PowerOk  Type: BOOL  Description: Voltage monitor readback |

|  |  |
| --- | --- |
| **Function Block**  FUNCTION\_BLOCK LowNoiseVcoFB  VAR\_INPUT  LowNoiseVcoType: LowNoiseVcoTypeEnum := VCO;  Request: SaveRestoreEnum;  LowNoiseVcoIn: LowNoiseVcoInStruct;  Frequency: LREAL := 0.0;  FrequencyError: BOOL := TRUE;  FddStages: INT := 1;  UseSigmaDelta: BOOL := TRUE;  END\_VAR  VAR\_OUTPUT  LowNoiseVcoOut: LowNoiseVcoOutStruct;  END\_VAR  VAR\_IN\_OUT  LowNoiseVcoInit: LowNoiseVcoStruct;  LowNoiseVco: LowNoiseVcoStruct;  END\_VAR | |
| Name | LowNoiseVcoFB |
| Description | Controls the low noise VCO. One function block for each low noise VCO chassis needs to be instantiated. An FDD unit is usually the second stage of a multi stage VCO/FDD setup. It does not implement a frequency servo. |
| Input argument | Name: LowNoiseVcoType  Type: LowNoiseVcoTypeEnum  Description: Type of low noise VCO chassis  Default: VCO |
| Input argument | Name: Request  Type: SaveRestoreEnum  Description: Save restore command |
| Input argument | Name: LowNoiseVcoIn  Type: LowNoiseVcoInStruct  Description: Input hardware structure |
| Input argument | Name: Frequency  Type: LREAL  Description: Externally measured frequency of VCO  Default: 0 |
| Input argument | Name: FrequencyError  Type: BOOL  Description: Externally measured frequency is invalid  Default: TRUE (invalid) |
| Input argument | Name: FddStages  Type: INT  Description: Number of frequency difference dividers used. This is to normalize the gain of the frequency servo.  Default: 1 |
| Input argument | Name: UseSigmaDelta  Type: BOOL  Description: Use a sigma delta modulator for averaging the control signal  Default: TRUE |
| Output argument | Name: LowNoiseVcoOut  Type: LowNoiseVcoOutStruct  Description: Output hardware structure |
| In/out argument | Name: LowNoiseVcoInit  Type: LowNoiseVcoStruct  Description: Save/restore variables in persistent memory |
| In/out argument | Name: LowNoiseVco  Type: LowNoiseVcoStruct  Description: User Interface structure |

|  |  |
| --- | --- |
| **Function Block**  FUNCTION\_BLOCK FrequencyDifferenceMixerFB  VAR\_INPUT  Request: SaveRestoreEnum;  FrequencyDifferenceMixerIn: LowNoiseVcoInStruct;  END\_VAR  VAR\_IN\_OUT  FrequencyDifferenceMixerInit: FrequencyDifferenceMixerStruct;  FrequencyDifferenceMixer: FrequencyDifferenceMixerStruct;  END\_VAR | |
| Name | FrequencyDifferenceMixerFB |
| Description | Controls the frequency difference mixer. One function block for each frequency difference mixer chassis needs to be instantiated. |
| Input argument | Name: Request  Type: SaveRestoreEnum  Description: Save restore command |
| Input argument | Name: FrequencyDifferenceMixerIn  Type: LowNoiseVcoInStruct  Description: Input hardware structure |
| In/out argument | Name: FrequencyDifferenceMixerInit  Type: FrequencyDifferenceMixerStruct  Description: Save/restore variables in persistent memory |
| In/out argument | Name: FrequencyDifferenceMixer  Type: FrequencyDifferenceMixerStruct  Description: User Interface structure |

|  |  |
| --- | --- |
| **Visual** | |
| Name | LowNoiseVcoVis |
| Description | Displays several MON and temperature readings, power and excitation status, and error alarms |
| Placeholder | Name: LowNoiseVCO  Type: LowNoiseVCOStruct  Description: Low Noise VCO structure |