Baseline Test Data for OMC In-vacuum DCPDs

R. Abbott, Caltech

30 July 2014

For Hardware Version: D060572-B2

- Overview

An analysis is performed on the DCPDs to compare them to the baseline results obtained from the prototype units manufactured for the eLIGO LLO OMC.

A full visual inspection is performed on the interior of each unit prior to welding shut

- DC Measurements:

  - DC Characteristics, quiescent current draw, no light.

    - + 15 volt supply current is 30mA, +/- 5mA

    - - 15 volt supply current is -30mA, +/- 5mA

  - Output DC Offsets

    - The DC offset voltages were measured differentially at the output of the pre-amplifier using a Fluke DC voltmeter

Table 1 (Output Offset Data)

| DCPD Serial Number | Specified DC Offset | Measured DC Offset |
| --- | --- | --- |
| S1400357 | +/- 200uV | 0.1 mVDC |

  - PD Bias line transient voltage suppressor (TVS) breakdown voltage corresponding to 1mA TVS current must be measured for each unit.  This revision of the design has a higher breakdown voltage TVS (18V vs. 12V on the original revision).  This was done to permit higher DCPD biases to allow higher photocurrents.

    - Voltage for 1mA TVS current conducting through D6 is 20V +1VDC, -0VDC for all three units.

| Specified D6, 1mA conduction Voltage | Measured D6, 1mA conduction Voltage |
| --- | --- |
| 20VDC +1VDC, -0VDC 20.8 VDC | 1 mA |

  - Measure the functionality of the transimpedance selection relay, and check that the input impedance indeed is 100 ohms and 400 ohms.

- Relay functional check OK [pass]

- Transfer Function and Noise:

  - Verify the transfer function and noise spectra are identical to the baseline data posted on the Advance LIGO wiki (+/-0.1dB for the transfer function, and +/- 1dB for the noise).

  - Use an oscilloscope to check each unit for high frequency oscillations, and none were observed.

  - Measured the bias line input to DCPD differential output transfer function to be in agreement with the baseline data.

- Transfer Function OK [pass]

- Noise OK [pass]

- No observed oscillation [pass]

- Bias transfer function OK [pass]