*LIGO Laboratory / LIGO Scientific Collaboration*

LIGO-E1400310-v2 *LIGO* 14 May 2015

aLIGO Availability Estimate

Dennis Coyne

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

|  |  |
| --- | --- |
| **California Institute of Technology****LIGO Project** | **Massachusetts Institute of Technology****LIGO Project** |
| **LIGO Hanford Observatory** | **LIGO Livingston Observatory** |

http://www.ligo.caltech.edu/

# Introduction

This is a preliminary estimate of the availability (or uptime) of the Advanced LIGO (aLIGO) interferometer. Updates to this availability estimate will be made as the system FMEA and reliability assessment is completed, and as more operational experience is gained.

# Applicable Documents

|  |  |
| --- | --- |
| [T080160](https://dcc.ligo.org/LIGO-T080160/public) | D. Hoak, B. Bland, “Downtime Accounting in S5”, Jul 2008 |
| [T1300519](https://dcc.ligo.org/LIGO-T1300519) | D. Coyne, “Sparing Analysis”,  |
| [T010075](https://dcc.ligo.org/LIGO-T010075) | P. Fritschel et. al., “Advanced LIGO Systems Design” |

# Scope

The scope of this availability estimate is a single aLIGO interferometer (not an estimate of coincidence operation, yet).

This estimate does not (yet) address the impact of non-interferometer, observatory systems on the system availability. For example facility systems such as HVAC and vacuum equipment. The experience from iLIGO/eLIGO is that these systems do not significantly impact the interferometer uptime (see section 6.9 and, for example, [T080160](https://dcc.ligo.org/LIGO-T080160)).

# Previous Experience

Some of the experience for Initial and Enhanced LIGO (iLIGO and eLIGO), in particular the S5 and S6 science runs are potentially relevant to anticipated aLIGO operations. Good downtime statistics have been assembled for the S5 run (T080160), but appear to be mostly lacking for the S6 run. (It is worth noting that LIGO operations intends to implement better operational metrics tracking that have been in place in the past.)

# Terminology

The terminology for the categories of the causes for interferometer downtime (unavailability) used here are very similar to those used in T080160.

## Uptime

Uptime is time the interferometer is in science mode. Science segments of any length are counted.

## Calibration

Calibration is time spent performing calibration studies. These are done periodically throughout the science runs. The same percentage time for iLIGO/eLIGO should apply to aLIGO.

## Hardware Failures

Hardware failures result in downtime due to needed repair or replacement of equipment. Examples are broken seismometers and electronics modules. The repair time includes diagnosis and isolation of the faulty equipment and the subsequent time to repair or replace and verify function (including reboots to computer equipment if appropriate), as well as the time required to bring the interferometer system back to low noise operation again.

Also included in this category is power grid outages.

## Software Failures

Software failures result in downtime due to faults in the software code. Examples are memory leaks and inadvertent triggers of software watchdogs (i.e. when not appropriate due to environmental conditions or equipment faults). The repair time includes diagnosis and re-start/re-boot of processes/computers as appropriate, as well as the time required to bring the interferometer system back to low noise operation again. The time required to isolate of the software fault, or define more appropriate parameter values (e.g. watchdog trigger levels), and then revise and regression test a new version of the code is assumed to be done off-line.

## Maintenance

Maintenance is downtime for planned detector maintenance, including scheduled repairs for failures which do not require prompt repair or replacement. During S5 & S6, four hours were scheduled every Tuesday for invasive, periodic maintenance tasks, such as liquid nitrogen deliveries, grounds-keeping, HEPI pump filter replacement, etc. This period is also used for scheduled detector maintenance such as PSL pump diode replacement, HWS SLED replacement, OptLev laser replacement, etc.

During iLIGO and eLIGO a 4 hour period per week was found to be sufficient. aLIGO is much more complex, so we expect that numerous diagnostic scripts will need to be run (as was done for HEPI at LLO for example) and more scheduled and deferred component repairs/replacements will be required in the weekly maintenance period. There presumably is a balance to be struck between diagnostic/maintenance time versus loss of run time due to failure that could have been avoided. Out engineering judgment is that we may require twice the time as for iLIGO (or 8 hr/week). In order to get a more precise estimate, we need an accounting from the Detector group of what tests need be done on what schedule. This should be possible to estimate after the first couple of engineering runs.

Observatory infrastructure will also need maintenance and repairs, some of which may bite into run time. During iLIGO and eLIGO Observatory maintenance did not reduce observing time. However our infrastructure is aging and so some small allocation seems reasonable.

## Commissioning

Commissioning is downtime for planned detector improvements. There were a few extended commissioning breaks at both sites during S5 and S6; these lasted for about a week and occurred once or twice a year. In addition, about 25 hours were allocated every month in S5 for discretionary studies.

## Transition to Low Noise Operation State

In the S5 downtime accounting document (T080160), the “scripts” category covered time spent running the "up" and "down" software scripts that take the instruments from lock-acquisition mode to the low noise detection mode. This required about ten minutes, each time the interferometer lost lock. The Automation System, Guardian, employed for aLIGO serves a similar function, but is more robust and should be more efficient, once the system is fully commissioned.

## Wind

The category “wind” includes downtime due to high winds. At Hanford, for S5 and S6 this is typically wind over 25-30 miles per hour.

In S5 and S6 at Livingston, the seismic pre-isolation system is very sensitive to the low-frequency ground motion caused by wind. Typically, Livingston could not operate in science mode when the wind was gusting more than 15mph. (At Livingston, this category also includes downtime due to Storms.)

## Seismic

For Hanford, the Seismic category includes any source of ground motion that prevents the interferometer from normal operations. Mostly, this downtime is due to earthquakes. The seismic environment is more complicated at Livingston, and this category is expanded into different sources:

Microseism is the term used to describe the low-frequency (0.1-0.35 Hz) ground motion caused by ocean waves striking the continental shelf. This is worst in the winter months, due to storms in the Atlantic.

Earthquakes produce large amplitude, low frequency ground motion (0.03-0.1 Hz), although higher frequencies can be driven depending on the earthquake's size and location. Typically, a 6.0-magnitude earthquake on the western Pacific Rim (Russia, Japan, the Philippines, or Indonesia) kept the Livingston interferometer out of lock for one or two hours in S5 and S6. Events like these happen about once every two days.

Trains are a phenomenon unique to Livingston. Railroad tracks pass by the site about 1.75 miles south of the ETMY end station, and freight trains move through Livingston three or four times a day. The largest trains occur regularly around 2am and around 7am, and typically take the interferometer out of low-noise operations for about half an hour.

# aLIGO Availability Estimate

## Uptime

The estimated uptime for a single aLIGO detector is 77%, as indicated in the Table and Figure below. The calibration period and the scheduled maintenance periods are scheduled to be simultaneous for the two observatories. The estimated dual coincidence uptime is 69%. Each of the factors that are part of this estimate are discussed in the sections below.





## Hardware Failures

The impact of availability due to hardware failures was estimated by taking the list of Line Replaceable/Repairable Units (LRUs) that was provided by the subsystem technical experts for each subsystem (and documented in T1300519, Sparing Analysis) and then projecting out the following three categories of failures:

* Repair/replacement of in-vacuum failed equipment [assigned to a once per year in-vacuum incursion event with a duration of 1 week duration, plus a 2 week pump-down period]
* Repair/replacement of in-air equipment which require prompt attention (i.e. failures which impact availability or performance, for which we have no credible early detection/forewarning of the failure)
* Repair/replacement of in-air equipment which can be scheduled (i.e. for which we have an early detection/forewarning of the failure) [assigned to be addressed in the weekly maintenance period]

For each of the failures, an estimate of the time required to diagnose, isolation, repair or replace and confirm function/operation has been made.

An estimate of the power outage rate for L1 was obtained by looking through the electronic logs for the period 6/1/2013 through 6/1/2014. Even short power “glitches” lasting a few seconds causes many of the interferometer detector systems to shut down and require a careful sequence of re-booting. For L1 there were 16 power outage events in the last year, of which only 5 had significant duration (many minutes or more).

At H1 there appear to have only been 4 events in the period 6/1/2013 through 6/1/2014. Note that in the entire past history at LHO a few (order of 3) power failures were due to onsite shorts rather than a PUD failure (e.g. a snake across the power bus).

It is currently difficult to estimate the duration of the impact of a power outage, beyond the duration of the outage itself. Here I have assumed a minimum 4 hour impact (in steady-state operations, once all procedures for re-start are well established). With this assumption, based on the last year of L1 experience, the total power outage impact is 83 hours, or 0.95%. This is the motivation to consider adding UPS systems[[1]](#footnote-1) to all interferometer detector systems (not just the mass storage room computers).





The 1 week vent, plus 2 week pump-down time, is for planned (scheduled) in-vacuum repairs which we find are are necessary/required, or desirable. These repairs/activities are defined in the table by reference to "event #1, in-vacuum maintenance". As noted in the table, we can operate with a single ISI sensor failure and a single OSEM failure per suspension stage (in general). (Note that these activities are in parallel, so one does not add the % times.) However, we can't operate with multiple sensor failures per assembly, so at some point we should plan a replacement. I put these activities into the maintenance category because it is deferred repair (not a prompt response to a hardware failure).

However, these estimates (from the subsystem experts) are based on some rather pessimistic MTBF estimates. While using pessimistic MTBF estimates are conservative for sparing analysis, they unrealistic skew the availability/uptime estimate. Some targeted re-evaluations of the MTBF estimates/guesses are necessary to improve the availability estimate. Note in particular that the pessimistic MTBF for the IO optics (coupled with their shear number) results in an unrealistically high "failure" (degradation) number per year.

The hardware failures are only for critical hardware. So PEM, HWS and OptLevs failures do not reduce downtime; They are assumed to be addressed as the opportunity arises during a weekly maintenance period. The table lists these scheduled repairs in the table by reference to "event #2, weekly maintenance period".

The items in the table with no event reference (NA) are assumed to require prompt attention and thus reduce availability directly:

## However note that repair of non-trivial hardware failures can only occur (more or less) immediately if they occur during normal work hours. If a non-trivial hardware failure occurs during the owl shift or a weekend, it will necessarily take considerably longer to repair, since it will require expert science or engineering staff to make the repair/replacement. This has not yet been factored into the analysis. This will increase somewhat the 3.48% estimate for repair time for failures requiring prompt attention.Software Failures

The rate of software failures in the steady-state after commissioning should be rather low. However given the complexity of the system as well as the need to update/maintain the software as operating systems and hardware platforms are updated, it is not reasonable to assume that there is no impact. For this estimate a 0.5% impact has been assumed until better information, or the basis for a better estimate, is developed.

## Maintenance

Just as for the S5 & S6 runs, a four hour maintenance period has been assumed weekly.

In order to accommodate the scheduled equipment maintenance/repair activities within this weekly 4-hour period, the observatory team would need to address ~3 repairs/tasks associated with the interferometer, on average, each week.

## Commissioning

In the aLIGO availability estimate I have assumed that no commissioning is being performed in steady-state operation.

## Transition to Low Noise Operation

The Automation System, Guardian, employed for aLIGO serves a similar function, but is more robust and should be more efficient, once the system is fully commissioned. The recent aLIGO lock acquisition for L1 using the Guardian takes about 10 minutes. Additional time will be required to bring the system to low noise operation as well. The goal is for lock acquisition lengths of > 40 hr. Here it has been assumed that the Guardian will require 15 minutes every 24 hr.

For the first aLIGO science run, the intent is to operate at 25W. When (eventually) aLIGO operates at high laser power, the time required to get to a low noise operational state, from a cold interferometer state, will be much longer. Even the interruption of high power operation for a short period will cause the interferometer optics to cool and require more time to re-establish the proper “hot” state. The availability estimate in this version does not take into account these considerations.

## Wind

The aLIGO seismic isolation system (SEI) was designed to provide much better seismic isolation performance at low frequencies than we had in iLIGO and eLIGO. Integrated test experience with aLIGO indicates that the promised SEI performance has basically been delivered. However the performance is a trade-off with robustness to environmental disturbances. This trade-off, and optimization for the environmental disturbances, has not yet been completed, and may not until we’re well into commissioning. There is evidence that wind at LHO causes significant ground tilt which the seismometer sensors sense as translation, and limits performance.

For this estimate, I have conservatively assumed that we don't improve on the sensitivity to wind-induced ground tilt sensitivity observed in S5 for LHO. In fact we are likely to improve our robustness to wind as the SEI systems are fully optimized.

## Seismic

As stated above, we fully expect (and have initial integrated test observations in support of this expectation) that we will have considerably better isolation from earthquakes and microseism disturbances. However we don’t yet have sufficient data to make assertions on the percentage of downtime with the new seismic isolation systems. For this estimate, I have assumed that downtime due to seismic disturbances is the same as H1 in S5 (and therefore considerably better for L1 than experienced in S5).

## Facility

### Vacuum Equipment

At LHO we used to have the LN dewar bumps which could cause lock loss this was fixed some time ago. LN deliveries sometimes continue outside our maintenance window. Other than the LN systems, the vacuum system has not caused any downtime. Ion pump failures are dealt with during the weekly maintenance period.

Of course past performance is not a guarantee of future performance. There is some evidence of vacuum system/component aging failures/issues, such as the leaks in two large gate valves at LLO. Repairs for the vacuum infrastructure may impact overall system availability, but this has not been factored into the availability estimate. Our baseline plan is to perform the vacuum repairs in parallel with commissioning activities between science runs.

### HVAC

During iLIGO/eLIGO chiller and fan failures resulted in a temperature variation during the time that we switch to the alternate unit. This resulted in reduced interferometer performance but no lock loss. Likewise storm fronts sometimes caused us to lose control of LVEA temperatures for a day or more with the resultant performance loss (not lock loss).

However the temperature sensitivity of the aLIGO system may be higher than for iLIGO/eLIGO. In particular, the aLIGO quad suspensions are far more sensitive to temperature. If the HVAC system is properly monitored/maintained, and the set point is appropriate, then HVAC is quite reliable and is unlikely to significantly effect overall system availability.

### Grounds keeping

Some nearby tumbleweed bailing at LHO has been problematic when next to a test mass (say within 100-200 feet) We can schedule this work to take place on the maintenance day - but sometimes there will be more than 4 hours of tumbleweeds or we need to gain access to an entry door. These are associated with wind events.

Storm fronts will sometimes cause us to lose control of LVEA temperatures for a day or more with the resultant performance loss (not lock loss).

1. A team began to define the requirements and potential design options for a facility-class UPS system. However the UPS team were asked to stand-down and document what work they've done to date. Reasons: (a) marginal case for UPS on the basis of down-time (estimated at 1% hit at LLO), (b) the LLO situation has improved; > 200 days without interruption, (c) we clearly do not have the staff to address the UPS system at the moment, (d) the installation and commissioning will cause significant schedule impact to ongoing commissioning. [↑](#footnote-ref-1)