## Slide 1: Noises in Gravitational Wave Detectors
Koji Arai – LIGO Laboratory / Caltech
LIGO-G1401145-v2

## Slide 2: Introduction ~ Noise?
GW detection Data stream of differential arm strain
Once recorded: Signals and noises are indistinguishable What we can do is to catch “likely” features
Reduce any kind of noises!

## Slide 3: Introduction ~ Noise?
Time domain vs frequency domain
Noise
Reduction
Time domain: transient noises Frequency domain: stationary noises

## Slide 4: Introduction ~ Noise?
Power Spectral Density (PSD) Double sided PSD (-Infinity < f < Infinity)
Single sided PSD (0 <= f < Infinity) [xunit2 / Hz]
Linearized PSD: [xunit/ sqrtHz]

## Slide 5: Introduction ~ Noise?
Parseval’s Theorem for signal RMS and PSD
Root Mean of x(t): average signal power density (per sec)
(cf. variance, std deviation)
PSD Sx(f):
power density per frequency (per sec)

## Slide 6: Introduction ~ Noise?
Example
PSD [fm/sqrtHz] in log-log scale, RMS [fm] ~ 50fm = 0.05pm
PSD [fm/sqrtHz] in log-lin scale
RMS [fm]
Time series [pm]

## Slide 7: Components of the interferometer
3 fundamentals of the GW detector
Mechanics
Optics
Electronics

## Slide 8: Components of the interferometer
3 fundamentals of the GW detector
Mechanics
Optics
Electronics

## Slide 9: Noise categories
3 fundamentals of the GW detector
Mechanics -> Displacement noises
Optics -> Optical noises
Electronics -> Electrical noises
Optical
Optical
Displacement
Displacement
Electrical
Electrical
Electrical

## Slide 10: Noise budget
Laser shot noise
Laser radiation
pressure noise
thermal noise
seismic noise
Laser intensity
/frequency noise
electronics noise
digitization noise
angular control noise
......

## Slide 11: Displacement noises

## Slide 12: Displacement noise
Mechanical displacement sensed by a laser interferometer
The longer the arm length, the smaller the strain noise
  Seismic noise
  Thermal noise
  Newtonian Gravity noise

## Slide 13: Displacement noise
Seismic noise
  Even when there is no noticeable earth quake…
http://link.aps.org/doi/10.1103/RevModPhys.86.121 (http://arxiv.org/abs/1305.5188)
Target disp. noise
10-20m/rtHz

## Slide 14: Displacement noise
Vibration isolation ~ utilize a harmonic oscillator
  A harmonic oscillator provides vibration isolation above its resonant frequency
w0=1Hz
|x/X|~1/1000@30Hz

## Slide 15: Displacement noise
How to get more isolation?
Damping
Lower the peak height
Worse isolation
Lower resonant freq
Better isolation
Complex to realize
Multi stage
Steeper isolation curve
More peaks
f -2
f -4
In practice: employ combination of these measures

## Slide 16: Displacement noise
iLIGO vibration isolation
Hydraulic active isolation / Isolation stack / Single Pendulum

## Slide 17: Displacement noise
aLIGO vibration isolation
Hydraulic active isolation / Invacuum Active Isolation Platforms / Multiple Pendulum

## Slide 18: Displacement noise
Virgo: super attenuator
  8m high
  9 stages in horizontal
  6 stages in vertical
8m

## Slide 19: Displacement noise
http://link.aps.org/doi/10.1103/RevModPhys.86.121 (http://arxiv.org/abs/1305.5188)

## Slide 20: Displacement noise
Question:
n-stage multiple pendulum with fixed height of H
How many stages n do we need to realize the vibration isolation of A at frequency of f?
For a given A what is the minimum f, we can realize by increasing n?
(Mass distribution)
  For equal m for each stage or
  For arbitrary mass mi and length hi

## Slide 21: Displacement noise
Thermal noise:
System in thermal equilibrium
  the system can dissipate its energy to the heat bath
  the system is thermally excited by thermal fluctuation
Mechanical thermal noises
  suspension thermal noise
  mirror substrate thermal noise
  mirror coating thermal noise

## Slide 22: Displacement noise
Fluctuation Dissipation Theorem
Friction: interaction with “bath” = huge number of d.o.f.
Fluctuation force: produced by huge number of d.o.f.
Dissipation and fluctuation have certain relationship
system description (Langevin equation)
q: generalized coordinate m: generalized mass
R: friction (dissipation)
F: internal force (restoring force, etc)
F’(t): fluctuating force from heat bath
Power spectrum density (PSD) of the fluctuation force

## Slide 23: Displacement noise
Transfer function approach
Equivalently, the fluctuation of the system can be obtained from the response of the system
where Z(w) and H(w) are the impedance and
force-to-displacement transfer function of the system

## Slide 24: Displacement noise
Question:
- Velocity damping of a pendulum
- Structure damping
loss angle: 0 < 𝜙 ≪ 1
- How anti-spring changes the thermal noise spectrum?
anti-spring parameter: 0<𝛼<1

## Slide 25: Displacement noise
In some cases, calculating the system response is complicated (e.g. deformation of an elastic body)
Systems response (impedance) at a certain freq:
Average rate of energy disspation
We don’t need to
know Z(w)!

## Slide 26: Displacement noise
Sensing of the mirror surface deformation with a laser beam (with intensity profile of f(r))
Apply periodic pressure with profile of f(r)
This induces deformation of x(r) which is different
from our sensing profile of f(r), but that’s OK
Calculate the rate of dissipation Wdiss
analytically, using FEA, or etc
Put this into the formula
We don’t need to
know Z(w)!
Y. Levin PRD 57, 659-663 (1998)

## Slide 27: Displacement noise
Mirror substrate thermal noise
  Brownian motion Mechanical loss associated with the internal friction Thermally excited body modes Optical coating (high mechanical loss) will be limiting noise source in aLIGO
  Thermo elastic noise Elastic strain & thermal expansion coefficient => cause heat distribution & flow in the substrate Temperature fluctuation causes mirror displacement
  Thermo-refractive noise Temp. fluctuation causes fluctuation of refractive index
We don’t need to
know Z(w)!

## Slide 28: Displacement noise
Suspension thermal noise
  Brownian motion Mechanical loss of the suspension fiber Thermally excited pendulum modes
  Thermo elastic noise Elastic strain of the fiber & thermal expansion coefficient => cause heat distribution & flow in the fiber Temperature fluctuation causes mirror motion
We don’t need to
know Z(w)!
<- Monolithic suspension
for high pendulum Q

## Slide 29: Displacement noise
Question
  Induced current damping (electro-mechanical system) 1. How does the Q factor of the system depend on R? 2. How much is the thermal noise displacement of the mass? 3. How does the thermal noise of the resister shakes the mass? 4. How are the above questions with a capacitive coupling instead of the coil?
  Cold damping 1. If the resister is cooled, how does the thermal noise motion change? 2. Is the pendulum actually cooled? Down to what temperature? 3. How fast the pendulum recovers the original temperature once R is returned to the room temp.?
We don’t need to
know Z(w)!

## Slide 30: Displacement noise
Newtonian Gravity noise
  Mass density fluctuations around the test masses => test mass motion via gravitational coupling
  Dominant source of Newtonian noise = Seismic surface wave
  Mitigation 1) Going to quiet place (underground) 2) Feedforward subtraction 3) Passive reduction by shaping local topography
NN
aLIGO
sensitivity
J Driggers, et al, PRD 86, 102001 (2012)
J Harms, et al, Class. Quantum Grav. 31 185011 (2014)

## Slide 31: Displacement noise
Mechanical upconversion noise
  Large low frequency (f < 1Hz) motion => upconverted to 10~100Hz motion via nonliner processes
  Barkhausen noise => low freq mirror actuation cause BH noise and upconversion Select better magnet materials (e.g. SmCo)
https://www.nde-ed.org/

## Slide 32: Optical noises

## Slide 33: Optical noises
Noises that contaminate the readout signal
  Quantum noises (shot noise, radiation pressure noise)
  Laser technical noises (frequency/intensity noise)
  Modulation noises
  Scattered light noise
Optical
Optical

## Slide 34: Optical noises
Quantum noises: Shot noise
  Noise due to photon counting statistics
  N detected photon => standard deviation √N
  Increasing the incident power Pin, => The shot noise is increased by √Pin => The signal amplitude is increased by Pin
  In total, the signal-to-noise ratio is improved by

## Slide 35: Optical noises
Quantum noises: Shot noise
  Photon shot noise associated with photodetection
  Michelson interferometer
  Michelson response (@DC)
at the limit of d𝜙->0
  iDC : DC Photocurrent
  η : PD Quantum Efficiency
  ν : Optical Frequency
Shot-noise limit of the Michelson phase sensitivity
  1.3x10-20 1/sqrtHz
  @1W, 1064nm, 4km
Michelson
Strain Sensitivity

## Slide 36: Optical noises
Supplemental slide ~ Shot noise derivation

## Slide 37: Optical noises
Quantum noises ~ Radiation pressure noise
  Photon number fluctuation in the arm cavity => Fluctuation of the back action force
  Quantum noise of the input laser => Common noise for two arms => cancelled and does not appear in the signal
  Vacuum fluctuation injected from the dark port => Differentially power fluctuation => Cause the noise in the GW signal

## Slide 38: Optical noises
Quantum noises
  Standard Quantum Limit (SQL)
- Trade-off Between Shot Noise and Radiation-Pressure Noise
- Uncertainty of the test mass position due to observation

## Slide 39: Optical noises
Laser frequency noise
  Laser wavelength (λ = c / ν) = reference for the displacement measurement
  Optical phase 𝜙 = 2 pi ν L / c d𝜙 = 2 pi / c (L dν + ν dL) <= indistinguishable
  dL/L target 10-24 => dν = 10-24 x 300 THz (1064nmYAG laser) = 3 x 10-10 Hz/rtHz

## Slide 40: Optical noises
Laser frequency noise
  Target: dνeff = 3 x 10-10 Hz/rtHz
  Laser stability dν = 10~100 Hz/rtHz @100Hz
3-stage cascaded laser frequency stabilization
iLIGO
  Michelson’s differential sensitivity provides Frequency noise cancellation of 1/100~1/1000 “Common Mode Rejection”

## Slide 41: Optical noises
Laser intensity noise
  Relative Intensity Noise (RIN): dP/P
  Sensor output V = P x => dV = P dx + x dP <= indistinguishable
  Requirement: RIN = 10-9 1/√Hz xofs=10e-12 (DC Readout) => dx=1e-20 m/√Hz

## Slide 42: Optical noises
Laser intensity noise ~ intensity stabiliaztion
  Requirement: RIN = 10-9 1/√Hz
  2-stage cascaded intensity stabilization control
  Challenge: requires 300mA of photodetection
  Shot noise limited RIN
  In-vacuum 8-branch Photodiode array
P. Kwee et al, Optics Express 20 10617-10634 (2012)

## Slide 43: Optical noises
Modulation noises
  RF Residual Amplitude Modulation
  Modulation Oscillator Phase Noise
  Modulation Oscillator Amplitude Noise
Produce noise sidebands on the modulation sidebands
Mitigation
  For the GW signal: Use DC readout and eliminate them by an “output mode cleaner cavity”
J. Camp et al JOSA A 17 120-128 (2000)

## Slide 44: Optical noises
Scattered light noise
  Scattered light recouples to the interferometer beam with an arbitrary phase => causes amplitude and phase fluctuation
  Two effects: 1. Small motion regime: linear coupling of the phase fluctuation 2. Large motion regime: low freq large motion of the scattering object => upconversion via fringe wrapping
  Mitigation
    Reduce scattered light
    Vibration isolation of the scattering object

## Slide 45: Electrical noises

## Slide 46: Electrical noises
General rules for electrical noises
  Electrical noise in photo detection
  Digitization noise (ADC/DAC) / Aliasing
Control noise
Actuator noise

## Slide 47: General rules for electrical noises
Low noise amplification at the beginning
Give necessary gain as early as possible
Don’t attenuate (and amplify again)
Gain G1
Gain G2
input noise: v1
input noise: v2
output noise: v1 G1
output noise Vout:
Sqrt[(v1 G1 G2)2+(V2 G2) 2]
Input equivalent noise
Vout / (G1 G2)
= Sqrt[v12+(v2/G1) 2]
Signal
Lessons
  The input referred noise is determined by v1
  It won’t become better by the later stages
  If G1 is big enough, we can ignore the noise of later stages

## Slide 48: Noise in photodetectors
Photodiodes
  PIN photodiodes (InGaAs for near IR, Si for visible)
    Good linearity
    Low noise
    High Quantum Efficiency (>90%)
InGaAs Quadrant PD
(𝜙3mm)
InGaAs PD (𝜙1mm)
P-layer
I-layer
N-layer
Anode
Cathode
“Photodiode Amplifiers”, J. Graeme (McGrawHill 1995)
Photocurrent

## Slide 49: Noise in photodetectors
Photodetectors are the first electrical block of the control chains
  It is important to have low input-referred current noise
Photo detection
  AF (Audio Frequency 0~100kHz)
    Prenty of light (photocurrent ~mA) Not a big electrical issue
  RF (Radio Frequency 10~200MHz)
    Large diode aperture -> high RF noise Need careful consideration

## Slide 50: Noise in photodetectors
input referred noise current
Noise in photodiodes
  Photodiode equivalent circuit
    Shunt Capacitance RD (~100MΩ) Usually not a problem
    Junction Capacitance CD (1pF~1nF)
    Series Resistance RS (1Ω~100Ω)
  The diode aperture size needs to be ~mm => Cd tends to be big.
  2mm InGaAs PD: Rs~10Ω, Cd~100pF => i_Rs = 20 pA/sqrtHz @100MHz
  (equivalent to the shot noise of 1mA light ~ 1.3mW@1064nm)

## Slide 51: Analog/Digital interface
Restriction of signal digitization
  Voltage quantization: quantization noise => limited dynamic range => Requires whitening/dewhitening filters
  Temporally discrete sampling: aliasing problem => limited signal bandwidth => Requires anti-aliasing (AA) / anti-imaging (AI) filters
  Typical signal chain

## Slide 52: Digitization (Quantization) noise
Analog signals (~+/-10V) -> Digital signal
  Digitized to a discrete N bit integer number
  Quantization causes a white noise e.g. +/-10V 16bit => Δ = 0.3mV => Vn ~ 100 μV/sqrtHz cf. Input noise of a typical analog circuit 10nV/sqrtHz
http://www.analog.com/static/imported-files/tutorials/MT-229.pdf
quantization
quantization error

## Slide 53: Digitization (Quantization) noise
Whitening
  Amplify a signal in the freq band where the signal is weak
Dewhitening
  Amplify a signal in the freq band where the signal is weak

## Slide 54: Control induced noise
Noise couplings from auxiliary loops
  e.g. Angle control feedback -> noise injection to the GW channel
  Mitigation 1) Make the coupling smaller 2) Make the noise itself smaller 3) Limit the control bandwidth of the aux loop

## Slide 55: Actuator noise
Actuator noise appears in the GW signal as an external disturbance
  Mitigation 1) Make the noise itself smaller 2) Make the actuator response smaller
  We need to keep sufficient actuator strength for lock acquisition => Transition to a low-noise mode after achieving lock

## Slide 56: Summary

## Slide 57: Summary
Summary
  There are such large number of noises
  They are quite omnidisciplinary
  Even only one noise can ruin our GW detection
  GW detection will be achieved by
    Careful design / knowledge / experience
    Logical, but inspirational trouble shooting
  Noise “hunting”