|  |  |  |  |
| --- | --- | --- | --- |
| **ECR Title: Add a 45.5MHz phase-locked OCXO** | | | DCC No: E1500371-v1 |
| Date: 8/27/2015 |
| **Requester: Daniel Sigg** | **Impacted Subsystem(s): ICS** | |  |
| **Description of Proposed Change(s):** We propose to add a 45.5 MHz phase-locked OCXO to clean up both amplitude and phase noise. Custom frequency oscillators have a long lead time and we propose to build 2 PLO chassis per detector. | | | |
| **Reason for Change(s):** Compared to the original design we are using very large modulation indices. This makes us more sensitive to RF AM. Furthermore, the 45.5MHz is generated as the 5th harmonics of the 9.1MHz which results in poor amplitude and phase noise. | | | |
| **Estimated Cost:** $4500/ifo for PLO&OCXO; WAG $3000/ifo for chassis & slow readback electronics. | | | |
| **Schedule Impact Estimate:** can go on in parallel. | | | |
| **Nature of Change (check all that apply):**  **Safety**  **Correct Hardware**  **Correct Documentation** | | **Improve Hardware**  **Improve/Clarify Documentation**  **Change Interface**  **Change Requirement** | |
| **Importance:**  **Desirable for ease of use, maintenance, safety**  **Desirable for improved performance, reliability**  **Essential for performance, reliability**  **Essential for function**  **Essential for safety** | | **Urgency:**  **No urgency**  **Desirable by date/event: \_after O1 \_\_\_\_**  **Essential by date/event: \_\_\_\_\_\_\_\_\_\_\_\_\_\_**  **Immediately (ASAP)** | |
| **Impacted Hardware (select all that apply):**  **Repair/Modify. List part & SNs: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**  **Scrap & Replace. List part & SNs:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**  **Installed units? List IFO, part & SNs: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**  **Future units to be built** | | **Impacted Documentation** (list all dwgs, design reports, test reports, specifications, etc.):  CDS rack drawings, EtherCAT system drawing, RF distribution drawing | |
| **Disposition of the proposed change(s):**  The disposition of this proposed engineering change request is to be completed by Systems Engineering and indicated in the “Notes and Changes” metadata field in the DCC entry for this ECR. The typical dispositions are as follows:   * **Additional Information Required**: in which case the additional information requested is defined. The ECR requester then re-submits the ECR with the new information using the same DCC number for the ECR but with the next version number. * **Rejected**: in which case the reason(s) for the rejection are to be given * **Approved** * **Approved with Caveat(s)**: in which case the caveat(s) are listed * **TRB**: the ECR is referred to an ad-hoc Technical Review Board for further evaluation and recommendation. It is the System Engineer’s (or designee’s) responsibility to organize the TRB. The System Engineer (or designee) then makes a technical decision based on the TRB’s recommendation. Links to the TRB’s documentation (charge, memos, final report, etc.) are to be added to the “Related Documents” field for this ECR. * **CCB**: a change request for approval of additional funds or schedule impact is to be submitted to the Configuration Control Board. Links to the CCB’s documentation (CR, etc.) are to be added to the “Related Documents” field for this ECR.   **Concurrence by Project Management:**  Acknowledgement/acceptance/approval of the disposition is to be indicated by the electronic “signature” feature in the DCC entry for this ECR, by one the following personnel:   * Systems Scientist * Systems Engineer * Deputy Systems Engineer | | | |