*LIGO Laboratory / LIGO Scientific Collaboration*

LIGO-T1500519-v3 *LIGO* January 26, 2017

Photon Calibrator Maintenance and Calibration Plan

Travis Sadecki, Richard Savage

Distribution of this document:

LIGO Scientific Collaboration

This is an internal working note

of the LIGO Laboratory.

|  |  |
| --- | --- |
| **California Institute of Technology**  **LIGO Project** | **Massachusetts Institute of Technology**  **LIGO Project** |
| **LIGO Hanford Observatory** | **LIGO Livingston Observatory** |

http://www.ligo.caltech.edu/

# Introduction

The purpose of this document is to detail the proposed Photon Calibrator maintenance and calibration schedule for both LHO and LLO. Note that the proposed schedule is subject to some float due to constraints regarding scheduled maintenance day activities, stand-down time related to Observing Runs, and availability of personnel.

# Schedule table of Photon Calibrator Maintenance and Calibrations

|  |  |  |  |
| --- | --- | --- | --- |
| **Task** | **Frequency** | **Site** | **Notes** |
| Gold Standard (GS) calibration | Once per year | LHO | Send the GS to NIST for calibration. |
| Working Standard 1 (WS1) / GS and Checking Standard (CS) ratio measurements | Every 2 weeks | LHO | Bench top ratio measurements in the PCal lab. |
| Working Standard 2 (WS2) / GS ratio measurement | Twice per year (January and July) | LLO | LLO sends WS2 to LHO for lab ratio measurements. |
| End station calibration (WS/RX and WS/TX measurements) | Every 2 months (X end first Tuesday of even months Y end first Tuesday of odd months) | LHO and LLO | Independent end station calibration measurements at each site. |
| Alignment and power budget check of RX and TX modules. | Twice per year | LHO and LLO | Check powers and alignments at both end station RX and TX modules for possible issues. |

All other non-routine maintenance activities will be scheduled on an as-needed/target of opportunity basis.