## Slide 1: VOPO Suspension Final Design
LIGO MIT Lab
Matichard, Fabrice
Fernandez Galiana, Alvaro

## Slide 2: VOPO Suspension
Design based on HAM 6
Maximizing Bench Surface: 3.64 ft2
Holes availability (Dog Clamps)
“Shooting” Area (for the outcome beam)
In Vacuum RFPD Area
Space for Tip Tilt

## Slide 3: VOPO Suspension
Blade with 3-axial symmetry
Blades clamps form equilateral triangle
Triangle center close to “Geometric Center” (easy to balance)
Stops separated enough

## Slide 4: VOPO Suspension
Conclusions:
Questions:
  Is it enough space for the Tip Tilt?
  Do we need a Tip Tilt?
  Is it enough shooting space?

## Slide 5: Optical Layout
2 different layouts:
  O3 Squeezer model with filter cavity + RFPD in Vacuum
  O2 Squeezer model (early squeezing)

## Slide 6: Optical Layout
Solid Works Optical Layout
Mirror:
  Mass: 205.87 g
  Quantity: 11 / 18
Lens:
  Mass: 121.375 g
  Quantity: 7 / 9

## Slide 7: Optical Layout
Solid Works Optical Layout
Beam Dump:
  Mass: 164.61 g
  Quantity: 2 / 5
QPD:
  Mass: 142.98 g
  Quantity: 2 / 3

## Slide 8: Optical Layout
Solid Works Optical Layout
Polarizers:
  Mass: 206.71 g
  Quantity: 2+2 / 2+2
Wave Plate:
  Mass: 299.69 g
  Quantity: 1 / 2
NOTE: At this point the polarizers have been modeled as mirrors. A more realistic model has to be done to the two kinds of polarizers that will be used.

## Slide 9: Optical Layout
Solid Works Optical Layout
Fibers In:
  Mass: 252.01 g
  Quantity: 2 / 2
RFPD in Vacuum
  Mass: 270.41 g
  Quantity: 0 / 2
NOTE: At this point the
fiber In have been modeled
as mirrors with pico motors. A
more realistic model has to be done to
the two kinds of polarizers that will be used.
NOTE: At this point the
model used is for a 4” beam
height while the real one will be at 2.5”

## Slide 10: Optical Layout
Solid Works Optical Layout
Faraday Rotator:
  Mass:
    Base: 170.43 g
    Rotator: X g
  Quantity: 1 / 1
NOTE: At this point the
Faraday Rotator has been
modeled as just its base and
with a rough design that will
be improved once the Faraday
rotator model will be obtained.
Base (6061-T6 Al)
EOT Faraday Rotator

## Slide 11: Optical Layout
Solid Works Optical Layout
VOPO Cavity:
  Mass:: 1463 g
  Quantity: 1 / 1

## Slide 12: Optical Layout
O3 (Filter Cavity + In Vacuum RFPD)
  Option 1

## Slide 13: Optical Layout
O3 (Filter Cavity + In Vacuum RFPD)
  Option 1
    Problems

## Slide 14: Optical Layout
O3 (Filter Cavity + In Vacuum RFPD)
  Option 1
    Problems

## Slide 15: Optical Layout
O3 (Filter Cavity + In Vacuum RFPD)
  Option 2

## Slide 16: Optical Layout
O3 (Filter Cavity + In Vacuum RFPD)
  Option 2
    Problems

## Slide 17: Optical Layout
O2

## Slide 18: Optical Layout
CONCLUSIONS
  Injection Bench is large enough
QUESTIONS

## Slide 19: Injection Bench
Stiffener first design (dummy)
Mass: 15.982 kg
f = 267.77 Hz

## Slide 20: Injection Bench
Stiffener first design (dummy II)
Mass: 15.977 kg
f = 276.29 Hz

## Slide 21: Injection Bench: Mass Budget
O3 (Filter Cavity + RFPD)
NOTE II: In this calculation the Faraday Rotator is not considered, just its base, and the model for polarizers and fiber in is set as the one of a mirror mount
NOTE I: The weight of the Injection Bench has been set at 18kg but it may vary after the frequency FEA
Optics: 12 kg (2.1kg contingency)
Bench: 18 kg
Balancing Mass: 6kg
SUSPENDED MASS: 36 kg

## Slide 22: Injection Bench: Mass Budget
O2
NOTE II: In this calculation the Faraday Rotator is not considered, just its base, and the model for polarizers and fiber in is set as the one of a mirror mount
Optics: 8 kg (1.2kg contingency)
Bench: 18 kg
Balancing Mass: 10kg
NOTE I: The weight of the Injection Bench has been set at 18kg but it may vary after the frequency FEA
SUSPENDED MASS: 36 kg

## Slide 23: Injection Bench Assembly FEA
Optical components for FEA
Mirror:
  Mass: 205.87 g
  Mass FEA: 203 g
  Quantity: 11 / 18
Lens:
  Mass: 121.375 g
  Mass FEA: 120 g
  Quantity: 7 / 10

## Slide 24: Injection Bench Assembly FEA
Optical components for FEA
Beam Dump:
  Mass: 164.61 g
  Mass to FEA: 163 g
  Quantity: 2 / 5
QPD:
  Mass: 142.98 g
  Mass FEA: 143 g
  Quantity: 2 / 3

## Slide 25: Injection Bench Assembly FEA
Optical components for FEA
Wave Plate:
  Mass: 300 g
  Mass FEA: 297 g
  Quantity: 1 / 2
Polarizer:
  Mass: 206.71 g
  Mass FEA: 203 g
  Quantity: 2+2 / 2+2

## Slide 26: Injection Bench Assembly FEA
Optical components for FEA
Fiber In:
  Mass: 252.01 g
  Mass FEA: 191 g
  Quantity: 2 / 2
RFPD in Vacuum
  Mass: 270.41 g
  Mass: 279 g
  Quantity: 0 / 2

## Slide 27: Injection Bench Assembly FEA
Optical components for FEA
VOPO Cavity:
  Mass: 1463 g
  Mass: 1460 g
  Quantity: 1 / 1
Faraday Rotator:
  Mass:
    Base: 170.43 g
    Base FEA: 170 g
    Rotator: X g
  Quantity: 1 / 1

## Slide 28: Injection Bench Assembly FEA
Injection Bench Assembly for FEA
Global Contact: Bonded

## Slide 29: Injection Bench Assembly FEA
Injection Bench Assembly for FEA (no MASSES)
NOTE: No significant difference has been observed between compatible and incompatible meshing

## Slide 30: BLADE DESIGN: Material
Material: 440C SSTL (E = 210 GPa, Y = 1800 MPa, UTS = 1970 MPa, ρ = 7800 kg/m3, ν = 0.285)
Total Suspended Mass: m = 36 kg
Charge per blade: P = 117.72N (12 kg)
Constraints:
Factor of Safety: ≥ 33.3% (FoS ≥ 3)
Desired frequency: f ≈ 1.5 Hz
http://www.pennstainless.com/stainless-grades/400-series-stainless/440c-stainless-steel/
http://www.makeitfrom.com/material-properties/Hardened-440C-Stainless-Steel

## Slide 31: BLADE DESIGN: Dimensions
Design Parameters:
  Blade Width: a
  Blade Length: l
  Blade thickness: h
Equations (for triangular blades):
3 Parameters
2 Equations
1 DOF
l = 280 mm
NOTE: The choice of this length has been made after checking the optical layout

## Slide 32: BLADE DESIGN: Dimensions
Theoretical Results (BladeDesign.m) for m = 36kg, f = 1.5Hz and FoS = 3:
l = 280 mm
b = 80.12 mm
h = 2.0282 mm
Kzz = 1065.9 N/m
Tip deflection = 110.4 mm
NOTE: 80mm is less than what was previously used for the blade design (85 mm) so it fits widely in the designed bench.

## Slide 33: BLADE DESIGN: Dimensions
Theoretical Results (BladeDesign.m) for m = 36kg, f = 1.6 Hz
l = 280 mm
b = 85 mm
h = 2.0761 mm
FoS = 3.3346 (30%)
Kzz = 1212.8 N/m
NOTE: In order to increase the FoS and knowing that 280 mm length and 85 mm wide blade fits in the design, the desired frequency has been raised to 1.6 Hz.

## Slide 34: BLADE DESIGN: Dimensions
FINAL DIMENSIONS:
l = 280 mm
b = 85 mm
h = 2.1 mm
f = 1.5252 Hz
Kzz = 1255.1 N/m
FoS = 3
f = 1.6277 Hz
Kzz = 1255.1 N/m
FoS = 3.5473
m = 41kg
m = 36kg
NOTE: The expected performances are those shown for 36 kg (that already include some contingency). However, the suspended mass could be raised up to 41 kg keeping FoS ≥ 3

## Slide 35: BLADE DESIGN: FEA
First Analysis using SW Simulation
Second Analysis using ANSYS
m = 36 kg
Force: Normal to Surface

## Slide 36: BLADE DESIGN: FEA
Deformation
SW SIMULATION
ANSYS
Difference= 0.09%

## Slide 37: BLADE DESIGN: FEA
SW SIMULATION
ANSYS
Stress (von Misses)
Difference= 0.40%

## Slide 38: BLADE DESIGN: FEA
Strain (Equivalent)
SW SIMULATION
ANSYS

## Slide 39: BLADE DESIGN: FEA
Conclusions:
FoS significantly lower than calculated (FoS = 3.55) [-7.32%]
SW and ANSYS show similar results
Possible different Strain Definition
Calculated Tip Deflection Angle: 33.16°

## Slide 40: FLEXURE DESIGN: Properties
Material: Music Wire Steel (E = 154-201 GPa, Y = 1600-2000 Mpa)
Total Suspended Mass: m = 36 kg
Charge per wire: P = 117.12 N
Minimum Tensile Strengths (E1100187)
NOTE: T1500539 shows a 84.6% average relation between Y and UTS. Therefore, in further calculations Y is calculated as 80% of the Minimum Tensile Strength
T1500539

## Slide 41: FLEXURE DESIGN: Music Wire
Material: Music Wire Steel (E = 201 GPa, Y = 1710 Mpa)
Total Suspended Mass: m = 36 kg
Charge per wire: P = 117.72 N
Desired frequency: f < 1.5 Hz
FoS ≥ 33.3% (FoS ≥ 3)
Lf = 130mm
d = 0.41 mm
f = 1.3992 Hz
Kxx = 927.5 N/m
FoS = 1.5508
NOTE: Diameter of 0.41mm enables to re-use the Faraday Isolator clamps but FoS = 1.55 is too low for our purposes
Calculated using
d = 0.41 mm as reference
FlexureDesign.m

## Slide 42: FLEXURE DESIGN: Music Wire
Material: Music Wire Steel (E = 201 GPa, Y = 1600 Mpa)
Total Suspended Mass: m = 36 kg
Charge per wire: P = 117.72 N
Desired frequency: f < 1.5 Hz
FoS ≥ 33.3% (FoS ≥ 3)
Lf = 130mm
d = 0.61 mm
f = 1.4203 Hz
Kxx = 955.62 N/m
FoS = 2.9148
NOTE: Diameter of 0.61mm raise FoS up to almost 3, which is considered enough as the minimal guaranteed tensile strength has been considered
Calculated using
d = 0.61 mm as reference
NOTE: Wire clamps will have to be redefined
FlexureDesign.m

## Slide 43: FLEXURE DESIGN: Theory
Frequency calculation using pendulum theory

123.1869 mm
lz

## Slide 44: CLAMP DESIGN: Material & Preload
SCREW MATERIAL
316 SSTL + HELICOILS
http://www.ssina.com/download_a_file/fasteners.pdf
Y = 344.74 Mpa (50 ksi)
μ = 0.5
Tension = 2290.7 N
T = 8.8702 Nm (78.51 in-lb)
clampDesign.m FoS = 3
Note: No Socket Head Screw for this material neither Silver Plated option (McMaster & Holo-Krome)
Note II: Torque similar to the recommended for Generic Screws in T1100066 (75 in-lb)

## Slide 45: CLAMP DESIGN: Material & Preload
SCREW MATERIAL
18-8 SSTL + HELICOILS
http://www.nuttybolts.com/nutty-bolts/?page_id=146
http://www.mcmaster.com/#socket-head-cap-screws/=z2xdjy
https://www.boltdepot.com/fastener-information/materials-and-grades/bolt-grade-chart.aspx
Y = 448.16 Mpa (65 ksi)
μ = 0.5
Tension = 3066.9 N
T = 11.8757 Nm (105.11 in-lb)
clampDesign.m FoS = 3
Note I: Torque similar to the recommended for 18-8 Holo-Krome in T1100066 (100 in-lb)

## Slide 46: CLAMP DESIGN: Material & Preload
SCREW MATERIAL
17-4 PH SSTL + HELICOILS
http://www.deltafastener.com/17-4%20ph.html
http://www.mcmaster.com/#=z2x9xd
Y = 861.84 Mpa (125 ksi)
μ = 0.5 (conservative)
Tension = 5897.8 N
T = 22.8376 Nm (202.13 in-lb)
clampDesign.m FoS = 3
Note: No Socket Head Screw for this material neither Silver Plated option (McMaster & Holo-Krome)
Note II: Torque similar to the recommended for Carbon Steel 1960 Series in T1100066 (200 in-lb)

## Slide 47: CLAMP DESIGN: Number of Screws
P
Pb
Preliminary calculation using ClampDesign.m
316 SSTL + HELICOILS
  Bolts needed: min 2 screw per row for FoS > 3
  FoS = 5.5859 (with 2 screw per row)
18-8 SSTL + HELICOILS
  Bolts needed: min 1 screw per row for FoS > 3
  FoS = 3.7393 (with 1 screw per row)
17-4 PH SSTL + HELICOILS
  Bolts needed: min 1 screw per row for FoS > 3
  FoS = 7.1909 (with 1 screw per row)
Note: The final number of screws per row and the material election will be driven by the contact calculation between the blade and the platform

## Slide 48: Balancing Mass
Preliminary Mass Balancing
  First Bench Design (≈17kg)
Note: The remaining 1kg has been reserved for the Bench Optimization and final CoG adjustment
Note II: This model is the one that has been used for the injection bench optimization and will be tunned according to the results of this optimization.

## Slide 49: Injection Bench Assembly FEA
Injection Bench Assembly for FEA with Masses
Global Contact: Bonded
Mesh: Compatible

## Slide 50: Injection Bench Assembly FEA
Injection Bench Assembly for FEA
NOTE: The remaining 940g will be used to position the CoG

## Slide 51: Injection Bench Assembly FEA
Injection Bench Assembly for FEA (ANSYS)
NOTE: The areas with high Strain should be reinforced using the 940g

## Slide 52: Injection Bench Assembly FEA
Injection Bench Assembly for FEA (no MASSES)
NOTE: The remaining 940g will be used to position the CoG and to improve the high strain areas

## Slide 53: Center of Mass
Optics O3
Desired Center of Mass
Optics CoM
Note: The Center of Mass of the optics is already close to the desired CoM (the barycenter of the clamps attachment points)

## Slide 54: Balancing Mass
Preliminary Mass Balancing
  First Bench Design (≈17kg)
  Lateral Masses
Note: The remaining 3.44kg (940g in the injection bench and the rest as balancing masses on the bench) should be placed near the dashed area in order to balance the assembly

## Slide 55: Balancing Mass
Optics + Balancing Mass
Desired Center of Mass
Real CoM
Note: The balancing masses are more distributed in the left part of the assembly. That could be solved by “naturally” placing the CoG of the Injection Bench opposed to the Optics one.

## Slide 56: Center of Mass
FEA

## Slide 57: Center of Mass
With balancing masses
Center of Mass
Geometric Center

## Slide 58: Center of Mass
Optics
Desired Center of Mass
Optics CoM
Note: The Center of Mass of the optics is already close to the desired CoM (the barycenter of the clamps attachment points)

## Slide 59: Balancing Mass
Preliminary Mass Balancing
  First Bench Design (≈16kg)
Note: The weight of the suspension (basically the Injection Bench) is approximate and will be defined after the optimization to increase the frequency

## Slide 60: Balancing Mass
Optics + Balancing Mass
Desired Center of Mass
Real CoM
Note: The balancing masses are basically placed in the lateral of the Injection Bench (5.4 kg). However, 2.6 kg are placed on the Bench as contingency for the optics weight

## Slide 61: Center of Mass
FEA

## Slide 62: Center of Mass
FEA
  First Frequency Study

## Slide 63: Center of Mass
Optics
Desired Center of Mass
Optics CoM
Note: The Center of Mass of the optics far from the desired but there is a lot of available space to locate the balancing masses

## Slide 64: Balancing Mass
Preliminary Mass Balancing
  First Bench Design (≈18kg)
Note: The weight of the suspension (basically the Injection Bench) is approximate and will be defined after the optimization to increase the frequency

## Slide 65: Balancing Mass
Optics + Balancing Mass
Desired Center of Mass
Real CoM
Note: In this case there is more balancing mass placed on the bench than in the laterals as it has to compensate the lack of optics

## Slide 66: Center of Mass
FEA

## Slide 67: Center of Mass
FEA
  First Frequency Study

## Slide 68: Optimization
FEA
  Bench Optimization for O3
Bench Mass = 17.296 kg

## Slide 69: Fixer
Screw #10-24 Static Calculation
  Load = P = 117.72N
  FoS = 32.1622
Maximum Play
  Vertical:
  Horizontal:
  Lateral: