## Slide 1
Michelle Stephens, Nathan Tomlin, Chris Yung, Ivan Ryger, Malcolm White, John Lehman
March 15, 2019
Toward the Next Generation of Standards for Laser Power Measurements

## Slide 2: Electrical Substitution Radiometer (ESR)
Absorber = cavity
electrical heater
Temperature
sensor (T)
thermal bath (T0)
incident radiation
Heat Link (W/K)

## Slide 3: Existing CW Laser Power Standards (C-series)
Absorber = cavity
electrical heater
Temperature
sensor (T)
thermal bath (T0)
incident radiation
conductance (G)
5 cm

## Slide 4: The Next Generation of Standards
Si substrate
W heater around perimeter
Vertically aligned carbon nanotube (VACNT) absorber
Superconducting or W
Transition edge sensor (TES) or commercial off-the-shelf (COTS) thermistor or thin-film VOx for temperature measurement
Cryogenic
W Heater
Absorber
Heat Link
Transition Edge Sensor
5 mm
Credit: Nathan Tomlin/ Chris Yung
Room Temperature
Heat Link
Absorber
Thermistors
Heater
20 mm
Compact Spectral Irradiance Monitor
1 mm
5 mm
Absorber
Heater
COTS thermistors
Link
Compact Total Irradiance Monitor
Absorber
10 mm
Heater
Link
COTS thermistor
Space Applications
Cryogenic Optical Fiber Power Meter
W Heater
Absorber
3 mm
Heat Link
Transition Edge Sensor
Fast Far Infrared Bolometer

## Slide 5: Cryogenic Optical Fiber Power Meter
Operation at 300 mW, 2 mm beam diameter
850 nm
1310, 1550 nm
Uncertainty 0.4%, k=2
  Dominated by optical fiber beamsplitter
Optical Fibers
Detectors at 4 K

## Slide 6: Performance Requirements
Absorber
20 mm
Link
Heater
Next-Gen C
Thermometers
Current standard
2 cm beam diameter
100 mW – 100 mW
325 nm – 1.93 mm
15 minute
All custom
0.6 – 0.9% accuracy, k = 2
2 cm beam diameter
100 mW – 100 mW
325 nm – 1.93 mm
2 minute
Mostly off-the-shelf
0.1% accuracy, k = 2

## Slide 7: Components of Next-Gen C Standard
Bolometer with Peltier Cooler and bridge circuit
TEC Controller
Lock-in
PID controller
Voltmeters
Vacuum

## Slide 8: Characterization
Bolometer time constant in vacuum, 23 s
(21 s predicted)
Heat link G in vacuum, 0.0081 W/K
Closed loop power measurement, 8 mW 633 nm

## Slide 9: Sensitivity to Room Temperature

## Slide 10: Position Dependent Response
Measured Laser Power vs. Position on Detector
Blue = Middle of Chip
Green = Bottom, 7 mm offset from middle, closer to heat link
Pink = Top, 7 mm offset from middle, farther from heat link
Credit: Florian Stuker

## Slide 11: Fast Far-Infrared Bolometer
VACNT
absorber
(3 mm dia.)
Nb TES
W heater
Electrical Substitution Fourier Transfer Spectroscopy
(collaboration with Solomon Woods, NIST, Gaithersburg)
Optical signal cancelled with heat pulse
(red-optical, green – electrical)
Measurements from 1 Hz to 50 Hz

## Slide 12: 300 mW, modulated
Free space cryogenic bolometer with heat link and heater configuration specifically designed for LIGO power and beam diameter
Brewster window
Closed loop operation for speed
Cryostat exists
Electronics for fast closed-loop operation are under development
Bolometer prototypes fabricated